## Slide 1: Measurements of Max-IV girder systems
IMMW 18
June 2013
Christopher W. Ostenfeld
Danfysik A/S, Denmark

## Slide 2: Outline
MAX IV magnet girders
  Description of Max IV girders
Hall probe field mapping
  Description of concept
  Stability of alignment
  Some measurements results
Harmonic coil measurements
  Description of concept
  Stability of the measurements
  Some measurement results

## Slide 3: Production of MAX IV magnet girders
M1
M2
U3
MAX IV designed for 3 GeV with very low emittance( multi-bend achromat)
Danfysik is producing 20 each of M1, M2 and U3 with 10-13 magnets
Designed by Maxlab*
In total 60 dipoles, 220 steerers, 160 quads, 120 sextupoles, 120 octupoles
680 magnets in total
Small Ø25 mm aperture in the multipoles
M1
U3
Girder positions in the ring achromats
*

## Slide 4: Girders as produced

## Slide 5: Top/bottom M1 girder as produced
Quadrupole
Octupole
Quadrupole
Octupole
Skew corrector
Normal corrector
Combined function
dipole with
pole face strips
Soft ends

## Slide 6: Measuring challenge
Harmonic coil measurement
  How to affix coil assembly
  Choice of bearings
Hall mapping
  Purchase of new 7 m Hall bench
  Build Hall probe( KSY 14) with temperature correction, and Hall supply
  Long measurement campaigns
Large volume of measurements
  680 individual sub segments in total
  Demands for data handling
Requirements for automatisation, to measure all of this efficiently

## Slide 7: Extensive tests for First Articles
Extended set of measurements
(coil and Hall)
Extended set of measurements
(coil and Hall)
Remove magnet
from bench
Extended set of measurements
(coil and Hall)
Remove top of magnet girder,
reinstall and remeasure
from bench

## Slide 8: Hall measurement requirements
Hall scans along the nominal trajectory, ± 15 mm(Dipole)
  Accomplished by scanning rectangular grid, followed by interpolation
  On-the fly( like insertion device mapping)
  Time for complete dipole map. Around 3 hours
Excitation curves(Dipole)
Excitation curves(Quadrupoles)
  Through available access holes
  S=0, x= ± 5 mm
Horizontal scans of quadrupoles
  X=±15 mm. S=-115,0,115

## Slide 9: Hall measurement of quadrupoles
Arrows show measurement lines

## Slide 10: Details on software for Hall bench
Fully automated measurements
  Magnet is powered up, and held for appropriate wait time
  Hall probe is zeroed in mu-metal shield prior to measurement
  Probe is homed on magnetic cone
  Measurement is carried out
    Dipole map, excitation, etc.
  Probe zero is checked in mu-metal after measurement, and data is corrected for drift
  Off-line correction for temperature
Voltage/NTC logged with Keithley 2701
Power supply control and motion control within Igor

## Slide 11: Hall probe mapper setup
Precision Hall mapper on long granite table
Laser feed back on longitudinal z-axis and linear encoders on x,y-axes
Alignment by scanning over magnetic pins, knife-edges
Short term position st.dev. < 2 μm, long term drift <10 μm
Alignment with magnet tip
 | Z (mm) | X (mm)
Round 1 | 2888.7707 | -61.7454
Round 2 | 2888.7697 | -61.7483
Round 3 | 2888.7695 | -61.7465
Round 4 | 2888.7698 | -61.7491
Round 5 | 2888.7709 | -61.7448
St. Dev. | 0.0006 | 0.0016
Short term repeatability
Scanning dipole from side & quadrupoles through small holes in yoke

## Slide 12: Hall probe field mapping
High stability temperature calibrated Hall probe allows on-the-fly mapping
Large field grid measured on-the-fly of the combined function dipoles
Interpolated on-the-fly data agree with step-and-go data within 1.6·10-4
Repeated measurements of field integral stable within 7·10-5
Very good alignment and stability
Chart: S [mm]
Field map on-the-fly in 1.5 h
Interpolated field values along trajectory

## Slide 13: Harmonic coil measurements
Several things to consider
  Single long coil vs short segmented coils
  Compensated vs non-compensated
  Integration vs voltage measurements
  Radial vs tangential coil
Other
  Support material
  Choice of bearings

## Slide 14: Harmonic coil measurements
Several things to consider
  Single long coil vs short segmented coils
  Compensated vs non-compensated
  Integration vs voltage measurements
  Radial vs tangential coil
Other
  Support material
  Choice of bearings

## Slide 15: Harmonic insertion coil concept
New concept ( For us, anyway……)
Coil inserted from girder end with external encoder and motor
  Precision Heidenhain encoder(scale drum) ERA 4202 C
Mechanical coil positioning from girder reference surfaces
15 o tangential coil with a 10.7 mm measurement radius
One short harmonic coil segment for each magnet
  Coil lengths between 150 mm and 250 mm
3-5 coil segments per support rod, 13 segments in total
Calibration of each segment → phase and gradient strength
Voltage integration using 1 channel Metrolab 5150 with Trigger module
Coil segment positions for long M1 & M2 ends

## Slide 16: Sub coil calibration
Phase and strength calibration
Due to differing coils lengths, several reference magnets are needed
For 150 mm segments, 100 mm EM sextupole is used
For 200-250 mm segments, temperature stabilized permanent magnet quadrupole is used
Magnet is rotated 180 degrees to find coil phase error relative to mechanical reference.
Compare with Hall/ stretch-wire results
Appropriate Gain/Phase chosen when measuring sub-magnets
Magnet | Segment | Gain | Phase [Grad]
QDE | 1 | 1.00113 | -65.507
OXY | 2 | 1.01146 | -65.516
QFE | 3 | 1.01083 | -19.868
OXX | 4 | 1.00818 | -65.429
COA (By/Bx) | 5 | 1.02909 | -19.892
Calibration data for sample 5 segment coil

## Slide 17: Details on software for coil bench
Running under Labview
Each coil segment is saved, with reference, to intrinsic phase and strength
User chooses appropriate coil to measure correct magnet sub segment.
Power supply control within Labview, and excitation cycle runs automatically
EXCEL templates developed
  Traceability to calibrations and coils used

## Slide 18: Stability of the harmonic measurements
Short term repeatability test
  Field gradient variation 0.4·10-4
  Higher harmonics below 0.1 unit
  Magnet rotation variation 0.01 mrad
  Very stable
Simple test of thermal stability
  Field gradient drift 0.2 unit/°C
  Magnet rotation drift -0.03 mrad/°C
  Low thermal drift
Long term stability, average for 6 quads
  Disassembly of test jig and yoke
  Three measurements over several days
  Field gradient variation 2·10-4
Test on M1 quad | Repeatability | Thermal change/°C | 
Field gradient strength | 0.4 | 0.2 | Unit, 10-4
Higher harmonics, n3-4 | < 0.1 | 0.2 | Unit, 10-4
Higher harmonics, n≥5 | < 0.03 | 0.2 | Unit, 10-4
Magnet center dX, dY | < 0.001 | < 0.001 | mm
Magnet rotation | 0.01 | -0.03 | mrad
Average result, 6 quads | Stability | 
Field gradient strength | 2 | Unit, 10-4
Higher harmonics, n3-6 | < 0.4 | Unit, 10-4
Higher harmonics, n≥7 | < 0.1 | Unit, 10-4
Magnet center dX,dY | 0.004 | mm
Magnet rotation | 0.14 | mrad
Short term repeatability test on a quad
Long term stability with jig/yoke disassembly

## Slide 19: Higher harmonics of a M2 quadrupole
Chart: Measured higher harmonic content and st.deviation at nominal current
| 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | 13 | 14 | 15 | 16 | 17 | 18
Amplitude dB/B0 [unit] | 4.4165 | 1.72 | 0.35381 | 9.7901 | 0.37836 | 0.10171 | 0.057575 | 1.4751 | 0.039053 | 0.043112 | 0.035549 | 0.78561 | 0.0111 | 0.033383 | 0.015632 | 0.14698
St.Dev. for amplitude [Unit] | 0.017169 | 0.029077 | 0.0097843 | 0.022652 | 0.011095 | 0.036754 | 0.014949 | 0.017984 | 0.0099079 | 0.027196 | 0.0063548 | 0.022564 | 0.0078304 | 0.0099591 | 0.011425 | 0.025043
Pole end contributions n=6,10,14 -> reduced for U1-5 quads by chamfering
Remaining harmonic errors: mainly sextupole (n=3) and octupole (n=4)
Remaining terms typically below 1 unit = 0.01%
Measuring noise level is low
Similar pattern for sextupole and octupole magnets
Sextupole
Octupole
Pole end contribution (n=6,10,14)
Low noise

## Slide 20: Next Project: Max IV 1.5 GeV
Same concept
12 identical double-bend achromats
96 m circumference
6 nm horizontal emittance
These magnets are much larger( around 4.5 m)
Exotic combined-function elements
Same exact order for Solaris in Krakow
Element | Description
SQFo | Combined function quadrupole/sextupole
SCo | Trim sextupole with skew quad, x- and y-corr windings
SDo | Sextupole
DIP | Gradient dipole with pole face strips
SDi | Sextupole
SCi | Trim sextupole with skew quad, x- and y-corr windings
SQFi | Combined function quadrupole/sextupole

## Slide 21: Still in design phase
4.5 m
Combined function
quadrupole/sextupole(SQFI)
Trim sextupole with
skew quad,
x- and y-corr windings
(Sci)
Challenge related to measuring
central multipoles, due to limited access

## Slide 22: Summary
We have devised a measurement scheme to measure Maxlab girders with high stability
Hall system shows high positional reproducibility, and stability
Succesful implementation of tangential, segmented coils
Harmonic coils can be reproducibly “lodged” in the girder
High degree of automatization implemented in Igor and Labview
Awaiting MaxLab 1.5 GeV girders for measurement

## Slide 23: Acknowledgements
Claus Hansen: Project Manager
Dan Kristoffersen: Project engineer
Franz Bødker: Physicist
Gunver Nielsen: Physicist
Esben Krauthammer: Physicist
Christian Glarbo: Physicist
Kristoffer Andersen: Measurement Technician