## Slide 1: Magnetic Measurements and Tuning of FLASH II Undulators.
P. Vagin*, P. Neumann, A. Schöps, M. Tischer, DESY, Germany
Presented by Markus Tischer
IMMW18, BNL, June 3rd-7th
* pavel.vagin@desy.de

## Slide 2: Conclusion
FLASH II - IDs
  Improved mechanical design based on PETRA IDs
  Smaller end-kick gap dependence
  Transfer measurements consider the actual pole contour
  Installation of 12 IDs in late 2013
Magnetic Measurements
  Hall probe calibration issues observed
  Characterization of various integrators in order to improve wire & coil measurements

## Slide 3: FLASH II

## Slide 4: FLASH II
HHG Seed laser
1st BC
2nd BC
FLASH II
FLASH II
Gun
ACC1+ACC39
ACC2&3
ACC4,5,6,7
Up to ~1.25 GeV
sFLASH
SASE
~4.1 – 44 nm
EXP. Hall
Beam parameters | 
Beam energy | 0.5 – 1.25 GeV
Normalized emittance | 1.4 – 3 mm mrad
Energy spread | 0.5 MeV
Peak current | 2.5 kA
Bunches per second | <8000
Bunch charge | 0.02 – 1 nC
SASE undulators | 
Period length | 31.4 mm
Segments length | 2.5 m
No. of segments | 12
Focusing structure | FODO
Radiation | SASE
Wavelength range SASE | 4-60 nm
FWHM pulse length | 10-500 fs
Peak power | 1-5 GW
Bandwidth | 0.5 – 2%
Number of pulses | <8000
Peak brilliance* | 1028-1031
Pulse energy | 1-500 µJ
* 1/(s mm2 mrad2 0.1%)

## Slide 5
FLASH II SASE Undulator Section
…
…
12 ID segments separated by diagnostic intersections
similar optics as for present sFLASH undulators
U32 undulators
  u = 31.4 mm
  gmin = 9.0 mm
  Bpeak = 0.98 T
  K = 2.87
  L = 12 × 2.5 m
Intersection comprises quadrupole, phase shifter (e.m.), BPM, OTR / wire scanner
Other IDs at FLASH2:
Optical modulator (1m)
THz em.-Undulator (3.5m)
Seeding IDs (future upgrade)
2nd harm. AfterBurner, (EPU,future)

## Slide 6: FLASH II undulator
Period: 31.4 mm
Gap: 9.0 mm
Field: 0.98 T
Kmax: 2.87
Length: 2.5 m
Load: 15 kN
Pole

## Slide 7: Support structure
1 µm
longitudinal axis
Vertical girder deformation
Same concept for FLASH II and PETRA Extension undulators (based on 2m devices used for PETRA III)
Various mechanical improvements (stiffness, adjustment, floor mount)
Less than 2 µm girder deformation under magnetic load
Revised drive system: 2 motors and 2 pairs of left/right-handed spindles
New linear encoder gap measurement system

## Slide 8: Transfer measurements
Touch probe measurement of
pole contour
Laser tracker measurement of girder fiducials
Matching touch probe and tracker data

## Slide 9: End structure
Parameters to optimize:
Last 2 magnets size (x,z), position(x,z)
Last 2 poles position(x)
Calculate endkick gap dependence of present configuration
Change each parameter a little
Calculate endkick gap dependence signature of each parameter
Find weights of signatures using least squares, that would compensate gap dependence best
Apply constrains: poles & magnets should not collide, limit maximum changes, to stay in linear range
Repeat

## Slide 10: Endkicks gap dependence
Chart: FLASH II trajectory
Chart: sFLASH, PETRAIII trajectory
| -1600 | -1599.5 | -1599 | -1598.5 | -1598 | -1597.5 | -1597 | -1596.5 | -1596 | -1595.5 | -1595 | -1594.5 | -1594 | -1593.5 | -1593 | -1592.5 | -1592 | -1591.5 | -1591 | -1590.5 | -1590 | -1589.5 | -1589 | -1588.5 | -1588 | -1587.5 | -1587 | -1586.5 | -1586 | -1585.5 | -1585 | -1584.5 | -1584 | -1583.5 | -1583 | -1582.5 | -1582 | -1581.5 | -1581 | -1580.5 | -1580 | -1579.5 | -1579 | -1578.5 | -1578 | -1577.5 | -1577 | -1576.5 | -1576 | -1575.5 | -1575 | -1574.5 | -1574 | -1573.5 | -1573 | -1572.5 | -1572 | -1571.5 | -1571 | -1570.5 | -1570 | -1569.5 | -1569 | -1568.5 | -1568 | -1567.5 | -1567 | -1566.5 | -1566 | -1565.5 | -1565 | -1564.5 | -1564 | -1563.5 | -1563 | -1562.5 | -1562 | -1561.5 | -1561 | -1560.5 | -1560 | -1559.5 | -1559 | -1558.5 | -1558 | -1557.5 | -1557 | -1556.5 | -1556 | -1555.5 | -1555 | -1554.5 | -1554 | -1553.5 | -1553 | -1552.5 | -1552 | -1551.5 | -1551 | -1550.5 | -1550 | -1549.5 | -1549 | -1548.5 | -1548 | -1547.5 | -1547 | -1546.5 | -1546 | -1545.5 | -1545 | -1544.5 | -1544 | -1543.5 | -1543 | -1542.5 | -1542 | -1541.5 | -1541 | -1540.5 | -1540 | -1539.5 | -1539 | -1538.5 | -1538 | -1537.5 | -1537 | -1536.5 | -1536 | -1535.5 | -1535 | -1534.5 | -1534 | -1533.5 | -1533 | -1532.5 | -1532 | -1531.5 | -1531 | -1530.5 | -1530 | -1529.5 | -1529 | -1528.5 | -1528 | -1527.5 | -1527 | -1526.5 | -1526 | -1525.5 | -1525 | -1524.5 | -1524 | -1523.5 | -1523 | -1522.5 | -1522 | -1521.5 | -1521 | -1520.5 | -1520 | -1519.5 | -1519 | -1518.5 | -1518 | -1517.5 | -1517 | -1516.5 | -1516 | -1515.5 | -1515 | -1514.5 | -1514 | -1513.5 | -1513 | -1512.5 | -1512 | -1511.5 | -1511 | -1510.5 | -1510 | -1509.5 | -1509 | -1508.5 | -1508 | -1507.5 | -1507 | -1506.5 | -1506 | -1505.5 | -1505 | -1504.5 | -1504 | -1503.5 | -1503 | -1502.5 | -1502 | -1501.5 | -1501 | -1500.5 | -1500 | -1499.5 | -1499 | -1498.5 | -1498 | -1497.5 | -1497 | -1496.5 | -1496 | -1495.5 | -1495 | -1494.5 | -1494 | -1493.5 | -1493 | -1492.5 | -1492 | -1491.5 | -1491 | -1490.5 | -1490 | -1489.5 | -1489 | -1488.5 | -1488 | -1487.5 | -1487 | -1486.5 | -1486 | -1485.5 | -1485 | -1484.5 | -1484 | -1483.5 | -1483 | -1482.5 | -1482 | -1481.5 | -1481 | -1480.5 | -1480 | -1479.5 | -1479 | -1478.5 | -1478 | -1477.5 | -1477 | -1476.5 | -1476 | -1475.5 | -1475 | -1474.5 | -1474 | -1473.5 | -1473 | -1472.5 | -1472 | -1471.5 | -1471 | -1470.5 | -1470 | -1469.5 | -1469 | -1468.5 | -1468 | -1467.5 | -1467 | -1466.5 | -1466 | -1465.5 | -1465 | -1464.5 | -1464 | -1463.5 | -1463 | -1462.5 | -1462 | -1461.5 | -1461 | -1460.5 | -1460 | -1459.5 | -1459 | -1458.5 | -1458 | -1457.5 | -1457 | -1456.5 | -1456 | -1455.5 | -1455 | -1454.5 | -1454 | -1453.5 | -1453 | -1452.5 | -1452 | -1451.5 | -1451 | -1450.5 | -1450 | -1449.5 | -1449 | -1448.5 | -1448 | -1447.5 | -1447 | -1446.5 | -1446 | -1445.5 | -1445 | -1444.5 | -1444 | -1443.5 | -1443 | -1442.5 | -1442 | -1441.5 | -1441 | -1440.5 | -1440 | -1439.5 | -1439 | -1438.5 | -1438 | -1437.5 | -1437 | -1436.5 | -1436 | -1435.5 | -1435 | -1434.5 | -1434 | -1433.5 | -1433 | -1432.5 | -1432 | -1431.5 | -1431 | -1430.5 | -1430 | -1429.5 | -1429 | -1428.5 | -1428 | -1427.5 | -1427 | -1426.5 | -1426 | -1425.5 | -1425 | -1424.5 | -1424 | -1423.5 | -1423 | -1422.5 | -1422 | -1421.5 | -1421 | -1420.5 | -1420 | -1419.5 | -1419 | -1418.5 | -1418 | -1417.5 | -1417 | -1416.5 | -1416 | -1415.5 | -1415 | -1414.5 | -1414 | -1413.5 | -1413 | -1412.5 | -1412 | -1411.5 | -1411 | -1410.5 | -1410 | -1409.5 | -1409 | -1408.5 | -1408 | -1407.5 | -1407 | -1406.5 | -1406 | -1405.5 | -1405 | -1404.5 | -1404 | -1403.5 | -1403 | -1402.5 | -1402 | -1401.5 | -1401 | -1400.5 | -1400 | -1399.5 | -1399 | -1398.5 | -1398 | -1397.5 | -1397 | -1396.5 | -1396 | -1395.5 | -1395 | -1394.5 | -1394 | -1393.5 | -1393 | -1392.5 | -1392 | -1391.5 | -1391 | -1390.5 | -1390 | -1389.5 | -1389 | -1388.5 | -1388 | -1387.5 | -1387 | -1386.5 | -1386 | -1385.5 | -1385 | -1384.5 | -1384 | -1383.5 | -1383 | -1382.5 | -1382 | -1381.5 | -1381 | -1380.5 | -1380 | -1379.5 | -1379 | -1378.5 | -1378 | -1377.5 | -1377 | -1376.5 | -1376 | -1375.5 | -1375 | -1374.5 | -1374 | -1373.5 | -1373 | -1372.5 | -1372 | -1371.5 | -1371 | -1370.5 | -1370 | -1369.5 | -1369 | -1368.5 | -1368 | -1367.5 | -1367 | -1366.5 | -1366 | -1365.5 | -1365 | -1364.5 | -1364 | -1363.5 | -1363 | -1362.5 | -1362 | -1361.5 | -1361 | -1360.5 | -1360 | -1359.5 | -1359 | -1358.5 | -1358 | -1357.5 | -1357 | -1356.5 | -1356 | -1355.5 | -1355 | -1354.5 | -1354 | -1353.5 | -1353 | -1352.5 | -1352 | -1351.5 | -1351 | -1350.5 | -1350 | -1349.5 | -1349 | -1348.5 | -1348 | -1347.5 | -1347 | -1346.5 | -1346 | -1345.5 | -1345 | -1344.5 | -1344 | -1343.5 | -1343 | -1342.5 | -1342 | -1341.5 | -1341 | -1340.5 | -1340 | -1339.5 | -1339 | -1338.5 | -1338 | -1337.5 | -1337 | -1336.5 | -1336 | -1335.5 | -1335 | -1334.5 | -1334 | -1333.5 | -1333 | -1332.5 | -1332 | -1331.5 | -1331 | -1330.5 | -1330 | -1329.5 | -1329 | -1328.5 | -1328 | -1327.5 | -1327 | -1326.5 | -1326 | -1325.5 | -1325 | -1324.5 | -1324 | -1323.5 | -1323 | -1322.5 | -1322 | -1321.5 | -1321 | -1320.5 | -1320 | -1319.5 | -1319 | -1318.5 | -1318 | -1317.5 | -1317 | -1316.5 | -1316 | -1315.5 | -1315 | -1314.5 | -1314 | -1313.5 | -1313 | -1312.5 | -1312 | -1311.5 | -1311 | -1310.5 | -1310 | -1309.5 | -1309 | -1308.5 | -1308 | -1307.5 | -1307 | -1306.5 | -1306 | -1305.5 | -1305 | -1304.5 | -1304 | -1303.5 | -1303 | -1302.5 | -1302 | -1301.5 | -1301 | -1300.5 | -1300 | -1299.5 | -1299 | -1298.5 | -1298 | -1297.5 | -1297 | -1296.5 | -1296 | -1295.5 | -1295 | -1294.5 | -1294 | -1293.5 | -1293 | -1292.5 | -1292 | -1291.5 | -1291 | -1290.5 | -1290 | -1289.5 | -1289 | -1288.5 | -1288 | -1287.5 | -1287 | -1286.5 | -1286 | -1285.5 | -1285 | -1284.5 | -1284 | -1283.5 | -1283 | -1282.5 | -1282 | -1281.5 | -1281 | -1280.5 | -1280 | -1279.5 | -1279 | -1278.5 | -1278 | -1277.5 | -1277 | -1276.5 | -1276 | -1275.5 | -1275 | -1274.5 | -1274 | -1273.5 | -1273 | -1272.5 | -1272 | -1271.5 | -1271 | -1270.5 | -1270 | -1269.5 | -1269 | -1268.5 | -1268 | -1267.5 | -1267 | -1266.5 | -1266 | -1265.5 | -1265 | -1264.5 | -1264 | -1263.5 | -1263 | -1262.5 | -1262 | -1261.5 | -1261 | -1260.5 | -1260 | -1259.5 | -1259 | -1258.5 | -1258 | -1257.5 | -1257 | -1256.5 | -1256 | -1255.5 | -1255 | -1254.5 | -1254 | -1253.5 | -1253 | -1252.5 | -1252 | -1251.5 | -1251 | -1250.5 | -1250 | -1249.5 | -1249 | -1248.5 | -1248 | -1247.5 | -1247 | -1246.5 | -1246 | -1245.5 | -1245 | -1244.5 | -1244 | -1243.5 | -1243 | -1242.5 | -1242 | -1241.5 | -1241 | -1240.5 | -1240 | -1239.5 | -1239 | -1238.5 | -1238 | -1237.5 | -1237 | -1236.5 | -1236 | -1235.5 | -1235 | -1234.5 | -1234 | -1233.5 | -1233 | -1232.5 | -1232 | -1231.5 | -1231 | -1230.5 | -1230 | -1229.5 | -1229 | -1228.5 | -1228 | -1227.5 | -1227 | -1226.5 | -1226 | -1225.5 | -1225 | -1224.5 | -1224 | -1223.5 | -1223 | -1222.5 | -1222 | -1221.5 | -1221 | -1220.5 | -1220 | -1219.5 | -1219 | -1218.5 | -1218 | -1217.5 | -1217 | -1216.5 | -1216 | -1215.5 | -1215 | -1214.5 | -1214 | -1213.5 | -1213 | -1212.5 | -1212 | -1211.5 | -1211 | -1210.5 | -1210 | -1209.5 | -1209 | -1208.5 | -1208 | -1207.5 | -1207 | -1206.5 | -1206 | -1205.5 | -1205 | -1204.5 | -1204 | -1203.5 | -1203 | -1202.5 | -1202 | -1201.5 | -1201 | -1200.5 | -1200 | -1199.5 | -1199 | -1198.5 | -1198 | -1197.5 | -1197 | -1196.5 | -1196 | -1195.5 | -1195 | -1194.5 | -1194 | -1193.5 | -1193 | -1192.5 | -1192 | -1191.5 | -1191 | -1190.5 | -1190 | -1189.5 | -1189 | -1188.5 | -1188 | -1187.5 | -1187 | -1186.5 | -1186 | -1185.5 | -1185 | -1184.5 | -1184 | -1183.5 | -1183 | -1182.5 | -1182 | -1181.5 | -1181 | -1180.5 | -1180 | -1179.5 | -1179 | -1178.5 | -1178 | -1177.5 | -1177 | -1176.5 | -1176 | -1175.5 | -1175 | -1174.5 | -1174 | -1173.5 | -1173 | -1172.5 | -1172 | -1171.5 | -1171 | -1170.5 | -1170 | -1169.5 | -1169 | -1168.5 | -1168 | -1167.5 | -1167 | -1166.5 | -1166 | -1165.5 | -1165 | -1164.5 | -1164 | -1163.5 | -1163 | -1162.5 | -1162 | -1161.5 | -1161 | -1160.5 | -1160 | -1159.5 | -1159 | -1158.5 | -1158 | -1157.5 | -1157 | -1156.5 | -1156 | -1155.5 | -1155 | -1154.5 | -1154 | -1153.5 | -1153 | -1152.5 | -1152 | -1151.5 | -1151 | -1150.5 | -1150 | -1149.5 | -1149 | -1148.5 | -1148 | -1147.5 | -1147 | -1146.5 | -1146 | -1145.5 | -1145 | -1144.5 | -1144 | -1143.5 | -1143 | -1142.5 | -1142 | -1141.5 | -1141 | -1140.5 | -1140 | -1139.5 | -1139 | -1138.5 | -1138 | -1137.5 | -1137 | -1136.5 | -1136 | -1135.5 | -1135 | -1134.5 | -1134 | -1133.5 | -1133 | -1132.5 | -1132 | -1131.5 | -1131 | -1130.5 | -1130 | -1129.5 | -1129 | -1128.5 | -1128 | -1127.5 | -1127 | -1126.5 | -1126 | -1125.5 | -1125 | -1124.5 | -1124 | -1123.5 | -1123 | -1122.5 | -1122 | -1121.5 | -1121 | -1120.5 | -1120 | -1119.5 | -1119 | -1118.5 | -1118 | -1117.5 | -1117 | -1116.5 | -1116 | -1115.5 | -1115 | -1114.5 | -1114 | -1113.5 | -1113 | -1112.5 | -1112 | -1111.5 | -1111 | -1110.5 | -1110 | -1109.5 | -1109 | -1108.5 | -1108 | -1107.5 | -1107 | -1106.5 | -1106 | -1105.5 | -1105 | -1104.5 | -1104 | -1103.5 | -1103 | -1102.5 | -1102 | -1101.5 | -1101 | -1100.5 | -1100 | -1099.5 | -1099 | -1098.5 | -1098 | -1097.5 | -1097 | -1096.5 | -1096 | -1095.5 | -1095 | -1094.5 | -1094 | -1093.5 | -1093 | -1092.5 | -1092 | -1091.5 | -1091 | -1090.5 | -1090 | -1089.5 | -1089 | -1088.5 | -1088 | -1087.5 | -1087 | -1086.5 | -1086 | -1085.5 | -1085 | -1084.5 | -1084 | -1083.5 | -1083 | -1082.5 | -1082 | -1081.5 | -1081 | -1080.5 | -1080 | -1079.5 | -1079 | -1078.5 | -1078 | -1077.5 | -1077 | -1076.5 | -1076 | -1075.5 | -1075 | -1074.5 | -1074 | -1073.5 | -1073 | -1072.5 | -1072 | -1071.5 | -1071 | -1070.5 | -1070 | -1069.5 | -1069 | -1068.5 | -1068 | -1067.5 | -1067 | -1066.5 | -1066 | -1065.5 | -1065 | -1064.5 | -1064 | -1063.5 | -1063 | -1062.5 | -1062 | -1061.5 | -1061 | -1060.5 | -1060 | -1059.5 | -1059 | -1058.5 | -1058 | -1057.5 | -1057 | -1056.5 | -1056 | -1055.5 | -1055 | -1054.5 | -1054 | -1053.5 | -1053 | -1052.5 | -1052 | -1051.5 | -1051 | -1050.5 | -1050 | -1049.5 | -1049 | -1048.5 | -1048 | -1047.5 | -1047 | -1046.5 | -1046 | -1045.5 | -1045 | -1044.5 | -1044 | -1043.5 | -1043 | -1042.5 | -1042 | -1041.5 | -1041 | -1040.5 | -1040 | -1039.5 | -1039 | -1038.5 | -1038 | -1037.5 | -1037 | -1036.5 | -1036 | -1035.5 | -1035 | -1034.5 | -1034 | -1033.5 | -1033 | -1032.5 | -1032 | -1031.5 | -1031 | -1030.5 | -1030 | -1029.5 | -1029 | -1028.5 | -1028 | -1027.5 | -1027 | -1026.5 | -1026 | -1025.5 | -1025 | -1024.5 | -1024 | -1023.5 | -1023 | -1022.5 | -1022 | -1021.5 | -1021 | -1020.5 | -1020 | -1019.5 | -1019 | -1018.5 | -1018 | -1017.5 | -1017 | -1016.5 | -1016 | -1015.5 | -1015 | -1014.5 | -1014 | -1013.5 | -1013 | -1012.5 | -1012 | -1011.5 | -1011 | -1010.5 | -1010 | -1009.5 | -1009 | -1008.5 | -1008 | -1007.5 | -1007 | -1006.5 | -1006 | -1005.5 | -1005 | -1004.5 | -1004 | -1003.5 | -1003 | -1002.5 | -1002 | -1001.5 | -1001 | -1000.5 | -1000 | -999.5 | -999 | -998.5 | -998 | -997.5 | -997 | -996.5 | -996 | -995.5 | -995 | -994.5 | -994 | -993.5 | -993 | -992.5 | -992 | -991.5 | -991 | -990.5 | -990 | -989.5 | -989 | -988.5 | -988 | -987.5 | -987 | -986.5 | -986 | -985.5 | -985 | -984.5 | -984 | -983.5 | -983 | -982.5 | -982 | -981.5 | -981 | -980.5 | -980 | -979.5 | -979 | -978.5 | -978 | -977.5 | -977 | -976.5 | -976 | -975.5 | -975 | -974.5 | -974 | -973.5 | -973 | -972.5 | -972 | -971.5 | -971 | -970.5 | -970 | -969.5 | -969 | -968.5 | -968 | -967.5 | -967 | -966.5 | -966 | -965.5 | -965 | -964.5 | -964 | -963.5 | -963 | -962.5 | -962 | -961.5 | -961 | -960.5 | -960 | -959.5 | -959 | -958.5 | -958 | -957.5 | -957 | -956.5 | -956 | -955.5 | -955 | -954.5 | -954 | -953.5 | -953 | -952.5 | -952 | -951.5 | -951 | -950.5 | -950 | -949.5 | -949 | -948.5 | -948 | -947.5 | -947 | -946.5 | -946 | -945.5 | -945 | -944.5 | -944 | -943.5 | -943 | -942.5 | -942 | -941.5 | -941 | -940.5 | -940 | -939.5 | -939 | -938.5 | -938 | -937.5 | -937 | -936.5 | -936 | -935.5 | -935 | -934.5 | -934 | -933.5 | -933 | -932.5 | -932 | -931.5 | -931 | -930.5 | -930 | -929.5 | -929 | -928.5 | -928 | -927.5 | -927 | -926.5 | -926 | -925.5 | -925 | -924.5 | -924 | -923.5 | -923 | -922.5 | -922 | -921.5 | -921 | -920.5 | -920 | -919.5 | -919 | -918.5 | -918 | -917.5 | -917 | -916.5 | -916 | -915.5 | -915 | -914.5 | -914 | -913.5 | -913 | -912.5 | -912 | -911.5 | -911 | -910.5 | -910 | -909.5 | -909 | -908.5 | -908 | -907.5 | -907 | -906.5 | -906 | -905.5 | -905 | -904.5 | -904 | -903.5 | -903 | -902.5 | -902 | -901.5 | -901 | -900.5 | -900 | -899.5 | -899 | -898.5 | -898 | -897.5 | -897 | -896.5 | -896 | -895.5 | -895 | -894.5 | -894 | -893.5 | -893 | -892.5 | -892 | -891.5 | -891 | -890.5 | -890 | -889.5 | -889 | -888.5 | -888 | -887.5 | -887 | -886.5 | -886 | -885.5 | -885 | -884.5 | -884 | -883.5 | -883 | -882.5 | -882 | -881.5 | -881 | -880.5 | -880 | -879.5 | -879 | -878.5 | -878 | -877.5 | -877 | -876.5 | -876 | -875.5 | -875 | -874.5 | -874 | -873.5 | -873 | -872.5 | -872 | -871.5 | -871 | -870.5 | -870 | -869.5 | -869 | -868.5 | -868 | -867.5 | -867 | -866.5 | -866 | -865.5 | -865 | -864.5 | -864 | -863.5 | -863 | -862.5 | -862 | -861.5 | -861 | -860.5 | -860 | -859.5 | -859 | -858.5 | -858 | -857.5 | -857 | -856.5 | -856 | -855.5 | -855 | -854.5 | -854 | -853.5 | -853 | -852.5 | -852 | -851.5 | -851 | -850.5 | -850 | -849.5 | -849 | -848.5 | -848 | -847.5 | -847 | -846.5 | -846 | -845.5 | -845 | -844.5 | -844 | -843.5 | -843 | -842.5 | -842 | -841.5 | -841 | -840.5 | -840 | -839.5 | -839 | -838.5 | -838 | -837.5 | -837 | -836.5 | -836 | -835.5 | -835 | -834.5 | -834 | -833.5 | -833 | -832.5 | -832 | -831.5 | -831 | -830.5 | -830 | -829.5 | -829 | -828.5 | -828 | -827.5 | -827 | -826.5 | -826 | -825.5 | -825 | -824.5 | -824 | -823.5 | -823 | -822.5 | -822 | -821.5 | -821 | -820.5 | -820 | -819.5 | -819 | -818.5 | -818 | -817.5 | -817 | -816.5 | -816 | -815.5 | -815 | -814.5 | -814 | -813.5 | -813 | -812.5 | -812 | -811.5 | -811 | -810.5 | -810 | -809.5 | -809 | -808.5 | -808 | -807.5 | -807 | -806.5 | -806 | -805.5 | -805 | -804.5 | -804 | -803.5 | -803 | -802.5 | -802 | -801.5 | -801 | -800.5 | -800 | -799.5 | -799 | -798.5 | -798 | -797.5 | -797 | -796.5 | -796 | -795.5 | -795 | -794.5 | -794 | -793.5 | -793 | -792.5 | -792 | -791.5 | -791 | -790.5 | -790 | -789.5 | -789 | -788.5 | -788 | -787.5 | -787 | -786.5 | -786 | -785.5 | -785 | -784.5 | -784 | -783.5 | -783 | -782.5 | -782 | -781.5 | -781 | -780.5 | -780 | -779.5 | -779 | -778.5 | -778 | -777.5 | -777 | -776.5 | -776 | -775.5 | -775 | -774.5 | -774 | -773.5 | -773 | -772.5 | -772 | -771.5 | -771 | -770.5 | -770 | -769.5 | -769 | -768.5 | -768 | -767.5 | -767 | -766.5 | -766 | -765.5 | -765 | -764.5 | -764 | -763.5 | -763 | -762.5 | -762 | -761.5 | -761 | -760.5 | -760 | -759.5 | -759 | -758.5 | -758 | -757.5 | -757 | -756.5 | -756 | -755.5 | -755 | -754.5 | -754 | -753.5 | -753 | -752.5 | -752 | -751.5 | -751 | -750.5 | -750 | -749.5 | -749 | -748.5 | -748 | -747.5 | -747 | -746.5 | -746 | -745.5 | -745 | -744.5 | -744 | -743.5 | -743 | -742.5 | -742 | -741.5 | -741 | -740.5 | -740 | -739.5 | -739 | -738.5 | -738 | -737.5 | -737 | -736.5 | -736 | -735.5 | -735 | -734.5 | -734 | -733.5 | -733 | -732.5 | -732 | -731.5 | -731 | -730.5 | -730 | -729.5 | -729 | -728.5 | -728 | -727.5 | -727 | -726.5 | -726 | -725.5 | -725 | -724.5 | -724 | -723.5 | -723 | -722.5 | -722 | -721.5 | -721 | -720.5 | -720 | -719.5 | -719 | -718.5 | -718 | -717.5 | -717 | -716.5 | -716 | -715.5 | -715 | -714.5 | -714 | -713.5 | -713 | -712.5 | -712 | -711.5 | -711 | -710.5 | -710 | -709.5 | -709 | -708.5 | -708 | -707.5 | -707 | -706.5 | -706 | -705.5 | -705 | -704.5 | -704 | -703.5 | -703 | -702.5 | -702 | -701.5 | -701 | -700.5 | -700 | -699.5 | -699 | -698.5 | -698 | -697.5 | -697 | -696.5 | -696 | -695.5 | -695 | -694.5 | -694 | -693.5 | -693 | -692.5 | -692 | -691.5 | -691 | -690.5 | -690 | -689.5 | -689 | -688.5 | -688 | -687.5 | -687 | -686.5 | -686 | -685.5 | -685 | -684.5 | -684 | -683.5 | -683 | -682.5 | -682 | -681.5 | -681 | -680.5 | -680 | -679.5 | -679 | -678.5 | -678 | -677.5 | -677 | -676.5 | -676 | -675.5 | -675 | -674.5 | -674 | -673.5 | -673 | -672.5 | -672 | -671.5 | -671 | -670.5 | -670 | -669.5 | -669 | -668.5 | -668 | -667.5 | -667 | -666.5 | -666 | -665.5 | -665 | -664.5 | -664 | -663.5 | -663 | -662.5 | -662 | -661.5 | -661 | -660.5 | -660 | -659.5 | -659 | -658.5 | -658 | -657.5 | -657 | -656.5 | -656 | -655.5 | -655 | -654.5 | -654 | -653.5 | -653 | -652.5 | -652 | -651.5 | -651 | -650.5 | -650 | -649.5 | -649 | -648.5 | -648 | -647.5 | -647 | -646.5 | -646 | -645.5 | -645 | -644.5 | -644 | -643.5 | -643 | -642.5 | -642 | -641.5 | -641 | -640.5 | -640 | -639.5 | -639 | -638.5 | -638 | -637.5 | -637 | -636.5 | -636 | -635.5 | -635 | -634.5 | -634 | -633.5 | -633 | -632.5 | -632 | -631.5 | -631 | -630.5 | -630 | -629.5 | -629 | -628.5 | -628 | -627.5 | -627 | -626.5 | -626 | -625.5 | -625 | -624.5 | -624 | -623.5 | -623 | -622.5 | -622 | -621.5 | -621 | -620.5 | -620 | -619.5 | -619 | -618.5 | -618 | -617.5 | -617 | -616.5 | -616 | -615.5 | -615 | -614.5 | -614 | -613.5 | -613 | -612.5 | -612 | -611.5 | -611 | -610.5 | -610 | -609.5 | -609 | -608.5 | -608 | -607.5 | -607 | -606.5 | -606 | -605.5 | -605 | -604.5 | -604 | -603.5 | -603 | -602.5 | -602 | -601.5 | -601 | -600.5 | -600 | -599.5 | -599 | -598.5 | -598 | -597.5 | -597 | -596.5 | -596 | -595.5 | -595 | -594.5 | -594 | -593.5 | -593 | -592.5 | -592 | -591.5 | -591 | -590.5 | -590 | -589.5 | -589 | -588.5 | -588 | -587.5 | -587 | -586.5 | -586 | -585.5 | -585 | -584.5 | -584 | -583.5 | -583 | -582.5 | -582 | -581.5 | -581 | -580.5 | -580 | -579.5 | -579 | -578.5 | -578 | -577.5 | -577 | -576.5 | -576 | -575.5 | -575 | -574.5 | -574 | -573.5 | -573 | -572.5 | -572 | -571.5 | -571 | -570.5 | -570 | -569.5 | -569 | -568.5 | -568 | -567.5 | -567 | -566.5 | -566 | -565.5 | -565 | -564.5 | -564 | -563.5 | -563 | -562.5 | -562 | -561.5 | -561 | -560.5 | -560 | -559.5 | -559 | -558.5 | -558 | -557.5 | -557 | -556.5 | -556 | -555.5 | -555 | -554.5 | -554 | -553.5 | -553 | -552.5 | -552 | -551.5 | -551 | -550.5 | -550 | -549.5 | -549 | -548.5 | -548 | -547.5 | -547 | -546.5 | -546 | -545.5 | -545 | -544.5 | -544 | -543.5 | -543 | -542.5 | -542 | -541.5 | -541 | -540.5 | -540 | -539.5 | -539 | -538.5 | -538 | -537.5 | -537 | -536.5 | -536 | -535.5 | -535 | -534.5 | -534 | -533.5 | -533 | -532.5 | -532 | -531.5 | -531 | -530.5 | -530 | -529.5 | -529 | -528.5 | -528 | -527.5 | -527 | -526.5 | -526 | -525.5 | -525 | -524.5 | -524 | -523.5 | -523 | -522.5 | -522 | -521.5 | -521 | -520.5 | -520 | -519.5 | -519 | -518.5 | -518 | -517.5 | -517 | -516.5 | -516 | -515.5 | -515 | -514.5 | -514 | -513.5 | -513 | -512.5 | -512 | -511.5 | -511 | -510.5 | -510 | -509.5 | -509 | -508.5 | -508 | -507.5 | -507 | -506.5 | -506 | -505.5 | -505 | -504.5 | -504 | -503.5 | -503 | -502.5 | -502 | -501.5 | -501 | -500.5 | -500 | -499.5 | -499 | -498.5 | -498 | -497.5 | -497 | -496.5 | -496 | -495.5 | -495 | -494.5 | -494 | -493.5 | -493 | -492.5 | -492 | -491.5 | -491 | -490.5 | -490 | -489.5 | -489 | -488.5 | -488 | -487.5 | -487 | -486.5 | -486 | -485.5 | -485 | -484.5 | -484 | -483.5 | -483 | -482.5 | -482 | -481.5 | -481 | -480.5 | -480 | -479.5 | -479 | -478.5 | -478 | -477.5 | -477 | -476.5 | -476 | -475.5 | -475 | -474.5 | -474 | -473.5 | -473 | -472.5 | -472 | -471.5 | -471 | -470.5 | -470 | -469.5 | -469 | -468.5 | -468 | -467.5 | -467 | -466.5 | -466 | -465.5 | -465 | -464.5 | -464 | -463.5 | -463 | -462.5 | -462 | -461.5 | -461 | -460.5 | -460 | -459.5 | -459 | -458.5 | -458 | -457.5 | -457 | -456.5 | -456 | -455.5 | -455 | -454.5 | -454 | -453.5 | -453 | -452.5 | -452 | -451.5 | -451 | -450.5 | -450 | -449.5 | -449 | -448.5 | -448 | -447.5 | -447 | -446.5 | -446 | -445.5 | -445 | -444.5 | -444 | -443.5 | -443 | -442.5 | -442 | -441.5 | -441 | -440.5 | -440 | -439.5 | -439 | -438.5 | -438 | -437.5 | -437 | -436.5 | -436 | -435.5 | -435 | -434.5 | -434 | -433.5 | -433 | -432.5 | -432 | -431.5 | -431 | -430.5 | -430 | -429.5 | -429 | -428.5 | -428 | -427.5 | -427 | -426.5 | -426 | -425.5 | -425 | -424.5 | -424 | -423.5 | -423 | -422.5 | -422 | -421.5 | -421 | -420.5 | -420 | -419.5 | -419 | -418.5 | -418 | -417.5 | -417 | -416.5 | -416 | -415.5 | -415 | -414.5 | -414 | -413.5 | -413 | -412.5 | -412 | -411.5 | -411 | -410.5 | -410 | -409.5 | -409 | -408.5 | -408 | -407.5 | -407 | -406.5 | -406 | -405.5 | -405 | -404.5 | -404 | -403.5 | -403 | -402.5 | -402 | -401.5 | -401 | -400.5 | -400 | -399.5 | -399 | -398.5 | -398 | -397.5 | -397 | -396.5 | -396 | -395.5 | -395 | -394.5 | -394 | -393.5 | -393 | -392.5 | -392 | -391.5 | -391 | -390.5 | -390 | -389.5 | -389 | -388.5 | -388 | -387.5 | -387 | -386.5 | -386 | -385.5 | -385 | -384.5 | -384 | -383.5 | -383 | -382.5 | -382 | -381.5 | -381 | -380.5 | -380 | -379.5 | -379 | -378.5 | -378 | -377.5 | -377 | -376.5 | -376 | -375.5 | -375 | -374.5 | -374 | -373.5 | -373 | -372.5 | -372 | -371.5 | -371 | -370.5 | -370 | -369.5 | -369 | -368.5 | -368 | -367.5 | -367 | -366.5 | -366 | -365.5 | -365 | -364.5 | -364 | -363.5 | -363 | -362.5 | -362 | -361.5 | -361 | -360.5 | -360 | -359.5 | -359 | -358.5 | -358 | -357.5 | -357 | -356.5 | -356 | -355.5 | -355 | -354.5 | -354 | -353.5 | -353 | -352.5 | -352 | -351.5 | -351 | -350.5 | -350 | -349.5 | -349 | -348.5 | -348 | -347.5 | -347 | -346.5 | -346 | -345.5 | -345 | -344.5 | -344 | -343.5 | -343 | -342.5 | -342 | -341.5 | -341 | -340.5 | -340 | -339.5 | -339 | -338.5 | -338 | -337.5 | -337 | -336.5 | -336 | -335.5 | -335 | -334.5 | -334 | -333.5 | -333 | -332.5 | -332 | -331.5 | -331 | -330.5 | -330 | -329.5 | -329 | -328.5 | -328 | -327.5 | -327 | -326.5 | -326 | -325.5 | -325 | -324.5 | -324 | -323.5 | -323 | -322.5 | -322 | -321.5 | -321 | -320.5 | -320 | -319.5 | -319 | -318.5 | -318 | -317.5 | -317 | -316.5 | -316 | -315.5 | -315 | -314.5 | -314 | -313.5 | -313 | -312.5 | -312 | -311.5 | -311 | -310.5 | -310 | -309.5 | -309 | -308.5 | -308 | -307.5 | -307 | -306.5 | -306 | -305.5 | -305 | -304.5 | -304 | -303.5 | -303 | -302.5 | -302 | -301.5 | -301 | -300.5 | -300 | -299.5 | -299 | -298.5 | -298 | -297.5 | -297 | -296.5 | -296 | -295.5 | -295 | -294.5 | -294 | -293.5 | -293 | -292.5 | -292 | -291.5 | -291 | -290.5 | -290 | -289.5 | -289 | -288.5 | -288 | -287.5 | -287 | -286.5 | -286 | -285.5 | -285 | -284.5 | -284 | -283.5 | -283 | -282.5 | -282 | -281.5 | -281 | -280.5 | -280 | -279.5 | -279 | -278.5 | -278 | -277.5 | -277 | -276.5 | -276 | -275.5 | -275 | -274.5 | -274 | -273.5 | -273 | -272.5 | -272 | -271.5 | -271 | -270.5 | -270 | -269.5 | -269 | -268.5 | -268 | -267.5 | -267 | -266.5 | -266 | -265.5 | -265 | -264.5 | -264 | -263.5 | -263 | -262.5 | -262 | -261.5 | -261 | -260.5 | -260 | -259.5 | -259 | -258.5 | -258 | -257.5 | -257 | -256.5 | -256 | -255.5 | -255 | -254.5 | -254 | -253.5 | -253 | -252.5 | -252 | -251.5 | -251 | -250.5 | -250 | -249.5 | -249 | -248.5 | -248 | -247.5 | -247 | -246.5 | -246 | -245.5 | -245 | -244.5 | -244 | -243.5 | -243 | -242.5 | -242 | -241.5 | -241 | -240.5 | -240 | -239.5 | -239 | -238.5 | -238 | -237.5 | -237 | -236.5 | -236 | -235.5 | -235 | -234.5 | -234 | -233.5 | -233 | -232.5 | -232 | -231.5 | -231 | -230.5 | -230 | -229.5 | -229 | -228.5 | -228 | -227.5 | -227 | -226.5 | -226 | -225.5 | -225 | -224.5 | -224 | -223.5 | -223 | -222.5 | -222 | -221.5 | -221 | -220.5 | -220 | -219.5 | -219 | -218.5 | -218 | -217.5 | -217 | -216.5 | -216 | -215.5 | -215 | -214.5 | -214 | -213.5 | -213 | -212.5 | -212 | -211.5 | -211 | -210.5 | -210 | -209.5 | -209 | -208.5 | -208 | -207.5 | -207 | -206.5 | -206 | -205.5 | -205 | -204.5 | -204 | -203.5 | -203 | -202.5 | -202 | -201.5 | -201 | -200.5 | -200 | -199.5 | -199 | -198.5 | -198 | -197.5 | -197 | -196.5 | -196 | -195.5 | -195 | -194.5 | -194 | -193.5 | -193 | -192.5 | -192 | -191.5 | -191 | -190.5 | -190 | -189.5 | -189 | -188.5 | -188 | -187.5 | -187 | -186.5 | -186 | -185.5 | -185 | -184.5 | -184 | -183.5 | -183 | -182.5 | -182 | -181.5 | -181 | -180.5 | -180 | -179.5 | -179 | -178.5 | -178 | -177.5 | -177 | -176.5 | -176 | -175.5 | -175 | -174.5 | -174 | -173.5 | -173 | -172.5 | -172 | -171.5 | -171 | -170.5 | -170 | -169.5 | -169 | -168.5 | -168 | -167.5 | -167 | -166.5 | -166 | -165.5 | -165 | -164.5 | -164 | -163.5 | -163 | -162.5 | -162 | -161.5 | -161 | -160.5 | -160 | -159.5 | -159 | -158.5 | -158 | -157.5 | -157 | -156.5 | -156 | -155.5 | -155 | -154.5 | -154 | -153.5 | -153 | -152.5 | -152 | -151.5 | -151 | -150.5 | -150 | -149.5 | -149 | -148.5 | -148 | -147.5 | -147 | -146.5 | -146 | -145.5 | -145 | -144.5 | -144 | -143.5 | -143 | -142.5 | -142 | -141.5 | -141 | -140.5 | -140 | -139.5 | -139 | -138.5 | -138 | -137.5 | -137 | -136.5 | -136 | -135.5 | -135 | -134.5 | -134 | -133.5 | -133 | -132.5 | -132 | -131.5 | -131 | -130.5 | -130 | -129.5 | -129 | -128.5 | -128 | -127.5 | -127 | -126.5 | -126 | -125.5 | -125 | -124.5 | -124 | -123.5 | -123 | -122.5 | -122 | -121.5 | -121 | -120.5 | -120 | -119.5 | -119 | -118.5 | -118 | -117.5 | -117 | -116.5 | -116 | -115.5 | -115 | -114.5 | -114 | -113.5 | -113 | -112.5 | -112 | -111.5 | -111 | -110.5 | -110 | -109.5 | -109 | -108.5 | -108 | -107.5 | -107 | -106.5 | -106 | -105.5 | -105 | -104.5 | -104 | -103.5 | -103 | -102.5 | -102 | -101.5 | -101 | -100.5 | -100 | -99.5 | -99 | -98.5 | -98 | -97.5 | -97 | -96.5 | -96 | -95.5 | -95 | -94.5 | -94 | -93.5 | -93 | -92.5 | -92 | -91.5 | -91 | -90.5 | -90 | -89.5 | -89 | -88.5 | -88 | -87.5 | -87 | -86.5 | -86 | -85.5 | -85 | -84.5 | -84 | -83.5 | -83 | -82.5 | -82 | -81.5 | -81 | -80.5 | -80 | -79.5 | -79 | -78.5 | -78 | -77.5 | -77 | -76.5 | -76 | -75.5 | -75 | -74.5 | -74 | -73.5 | -73 | -72.5 | -72 | -71.5 | -71 | -70.5 | -70 | -69.5 | -69 | -68.5 | -68 | -67.5 | -67 | -66.5 | -66 | -65.5 | -65 | -64.5 | -64 | -63.5 | -63 | -62.5 | -62 | -61.5 | -61 | -60.5 | -60 | -59.5 | -59 | -58.5 | -58 | -57.5 | -57 | -56.5 | -56 | -55.5 | -55 | -54.5 | -54 | -53.5 | -53 | -52.5 | -52 | -51.5 | -51 | -50.5 | -50 | -49.5 | -49 | -48.5 | -48 | -47.5 | -47 | -46.5 | -46 | -45.5 | -45 | -44.5 | -44 | -43.5 | -43 | -42.5 | -42 | -41.5 | -41 | -40.5 | -40 | -39.5 | -39 | -38.5 | -38 | -37.5 | -37 | -36.5 | -36 | -35.5 | -35 | -34.5 | -34 | -33.5 | -33 | -32.5 | -32 | -31.5 | -31 | -30.5 | -30 | -29.5 | -29 | -28.5 | -28 | -27.5 | -27 | -26.5 | -26 | -25.5 | -25 | -24.5 | -24 | -23.5 | -23 | -22.5 | -22 | -21.5 | -21 | -20.5 | -20 | -19.5 | -19 | -18.5 | -18 | -17.5 | -17 | -16.5 | -16 | -15.5 | -15 | -14.5 | -14 | -13.5 | -13 | -12.5 | -12 | -11.5 | -11 | -10.5 | -10 | -9.5 | -9 | -8.5 | -8 | -7.5 | -7 | -6.5 | -6 | -5.5 | -5 | -4.5 | -4 | -3.5 | -3 | -2.5 | -2 | -1.5 | -1 | -0.5 | 0 | 0.5 | 1 | 1.5 | 2 | 2.5 | 3 | 3.5 | 4 | 4.5 | 5 | 5.5 | 6 | 6.5 | 7 | 7.5 | 8 | 8.5 | 9 | 9.5 | 10 | 10.5 | 11 | 11.5 | 12 | 12.5 | 13 | 13.5 | 14 | 14.5 | 15 | 15.5 | 16 | 16.5 | 17 | 17.5 | 18 | 18.5 | 19 | 19.5 | 20 | 20.5 | 21 | 21.5 | 22 | 22.5 | 23 | 23.5 | 24 | 24.5 | 25 | 25.5 | 26 | 26.5 | 27 | 27.5 | 28 | 28.5 | 29 | 29.5 | 30 | 30.5 | 31 | 31.5 | 32 | 32.5 | 33 | 33.5 | 34 | 34.5 | 35 | 35.5 | 36 | 36.5 | 37 | 37.5 | 38 | 38.5 | 39 | 39.5 | 40 | 40.5 | 41 | 41.5 | 42 | 42.5 | 43 | 43.5 | 44 | 44.5 | 45 | 45.5 | 46 | 46.5 | 47 | 47.5 | 48 | 48.5 | 49 | 49.5 | 50 | 50.5 | 51 | 51.5 | 52 | 52.5 | 53 | 53.5 | 54 | 54.5 | 55 | 55.5 | 56 | 56.5 | 57 | 57.5 | 58 | 58.5 | 59 | 59.5 | 60 | 60.5 | 61 | 61.5 | 62 | 62.5 | 63 | 63.5 | 64 | 64.5 | 65 | 65.5 | 66 | 66.5 | 67 | 67.5 | 68 | 68.5 | 69 | 69.5 | 70 | 70.5 | 71 | 71.5 | 72 | 72.5 | 73 | 73.5 | 74 | 74.5 | 75 | 75.5 | 76 | 76.5 | 77 | 77.5 | 78 | 78.5 | 79 | 79.5 | 80 | 80.5 | 81 | 81.5 | 82 | 82.5 | 83 | 83.5 | 84 | 84.5 | 85 | 85.5 | 86 | 86.5 | 87 | 87.5 | 88 | 88.5 | 89 | 89.5 | 90 | 90.5 | 91 | 91.5 | 92 | 92.5 | 93 | 93.5 | 94 | 94.5 | 95 | 95.5 | 96 | 96.5 | 97 | 97.5 | 98 | 98.5 | 99 | 99.5 | 100 | 100.5 | 101 | 101.5 | 102 | 102.5 | 103 | 103.5 | 104 | 104.5 | 105 | 105.5 | 106 | 106.5 | 107 | 107.5 | 108 | 108.5 | 109 | 109.5 | 110 | 110.5 | 111 | 111.5 | 112 | 112.5 | 113 | 113.5 | 114 | 114.5 | 115 | 115.5 | 116 | 116.5 | 117 | 117.5 | 118 | 118.5 | 119 | 119.5 | 120 | 120.5 | 121 | 121.5 | 122 | 122.5 | 123 | 123.5 | 124 | 124.5 | 125 | 125.5 | 126 | 126.5 | 127 | 127.5 | 128 | 128.5 | 129 | 129.5 | 130 | 130.5 | 131 | 131.5 | 132 | 132.5 | 133 | 133.5 | 134 | 134.5 | 135 | 135.5 | 136 | 136.5 | 137 | 137.5 | 138 | 138.5 | 139 | 139.5 | 140 | 140.5 | 141 | 141.5 | 142 | 142.5 | 143 | 143.5 | 144 | 144.5 | 145 | 145.5 | 146 | 146.5 | 147 | 147.5 | 148 | 148.5 | 149 | 149.5 | 150 | 150.5 | 151 | 151.5 | 152 | 152.5 | 153 | 153.5 | 154 | 154.5 | 155 | 155.5 | 156 | 156.5 | 157 | 157.5 | 158 | 158.5 | 159 | 159.5 | 160 | 160.5 | 161 | 161.5 | 162 | 162.5 | 163 | 163.5 | 164 | 164.5 | 165 | 165.5 | 166 | 166.5 | 167 | 167.5 | 168 | 168.5 | 169 | 169.5 | 170 | 170.5 | 171 | 171.5 | 172 | 172.5 | 173 | 173.5 | 174 | 174.5 | 175 | 175.5 | 176 | 176.5 | 177 | 177.5 | 178 | 178.5 | 179 | 179.5 | 180 | 180.5 | 181 | 181.5 | 182 | 182.5 | 183 | 183.5 | 184 | 184.5 | 185 | 185.5 | 186 | 186.5 | 187 | 187.5 | 188 | 188.5 | 189 | 189.5 | 190 | 190.5 | 191 | 191.5 | 192 | 192.5 | 193 | 193.5 | 194 | 194.5 | 195 | 195.5 | 196 | 196.5 | 197 | 197.5 | 198 | 198.5 | 199 | 199.5 | 200 | 200.5 | 201 | 201.5 | 202 | 202.5 | 203 | 203.5 | 204 | 204.5 | 205 | 205.5 | 206 | 206.5 | 207 | 207.5 | 208 | 208.5 | 209 | 209.5 | 210 | 210.5 | 211 | 211.5 | 212 | 212.5 | 213 | 213.5 | 214 | 214.5 | 215 | 215.5 | 216 | 216.5 | 217 | 217.5 | 218 | 218.5 | 219 | 219.5 | 220 | 220.5 | 221 | 221.5 | 222 | 222.5 | 223 | 223.5 | 224 | 224.5 | 225 | 225.5 | 226 | 226.5 | 227 | 227.5 | 228 | 228.5 | 229 | 229.5 | 230 | 230.5 | 231 | 231.5 | 232 | 232.5 | 233 | 233.5 | 234 | 234.5 | 235 | 235.5 | 236 | 236.5 | 237 | 237.5 | 238 | 238.5 | 239 | 239.5 | 240 | 240.5 | 241 | 241.5 | 242 | 242.5 | 243 | 243.5 | 244 | 244.5 | 245 | 245.5 | 246 | 246.5 | 247 | 247.5 | 248 | 248.5 | 249 | 249.5 | 250 | 250.5 | 251 | 251.5 | 252 | 252.5 | 253 | 253.5 | 254 | 254.5 | 255 | 255.5 | 256 | 256.5 | 257 | 257.5 | 258 | 258.5 | 259 | 259.5 | 260 | 260.5 | 261 | 261.5 | 262 | 262.5 | 263 | 263.5 | 264 | 264.5 | 265 | 265.5 | 266 | 266.5 | 267 | 267.5 | 268 | 268.5 | 269 | 269.5 | 270 | 270.5 | 271 | 271.5 | 272 | 272.5 | 273 | 273.5 | 274 | 274.5 | 275 | 275.5 | 276 | 276.5 | 277 | 277.5 | 278 | 278.5 | 279 | 279.5 | 280 | 280.5 | 281 | 281.5 | 282 | 282.5 | 283 | 283.5 | 284 | 284.5 | 285 | 285.5 | 286 | 286.5 | 287 | 287.5 | 288 | 288.5 | 289 | 289.5 | 290 | 290.5 | 291 | 291.5 | 292 | 292.5 | 293 | 293.5 | 294 | 294.5 | 295 | 295.5 | 296 | 296.5 | 297 | 297.5 | 298 | 298.5 | 299 | 299.5 | 300 | 300.5 | 301 | 301.5 | 302 | 302.5 | 303 | 303.5 | 304 | 304.5 | 305 | 305.5 | 306 | 306.5 | 307 | 307.5 | 308 | 308.5 | 309 | 309.5 | 310 | 310.5 | 311 | 311.5 | 312 | 312.5 | 313 | 313.5 | 314 | 314.5 | 315 | 315.5 | 316 | 316.5 | 317 | 317.5 | 318 | 318.5 | 319 | 319.5 | 320 | 320.5 | 321 | 321.5 | 322 | 322.5 | 323 | 323.5 | 324 | 324.5 | 325 | 325.5 | 326 | 326.5 | 327 | 327.5 | 328 | 328.5 | 329 | 329.5 | 330 | 330.5 | 331 | 331.5 | 332 | 332.5 | 333 | 333.5 | 334 | 334.5 | 335 | 335.5 | 336 | 336.5 | 337 | 337.5 | 338 | 338.5 | 339 | 339.5 | 340 | 340.5 | 341 | 341.5 | 342 | 342.5 | 343 | 343.5 | 344 | 344.5 | 345 | 345.5 | 346 | 346.5 | 347 | 347.5 | 348 | 348.5 | 349 | 349.5 | 350 | 350.5 | 351 | 351.5 | 352 | 352.5 | 353 | 353.5 | 354 | 354.5 | 355 | 355.5 | 356 | 356.5 | 357 | 357.5 | 358 | 358.5 | 359 | 359.5 | 360 | 360.5 | 361 | 361.5 | 362 | 362.5 | 363 | 363.5 | 364 | 364.5 | 365 | 365.5 | 366 | 366.5 | 367 | 367.5 | 368 | 368.5 | 369 | 369.5 | 370 | 370.5 | 371 | 371.5 | 372 | 372.5 | 373 | 373.5 | 374 | 374.5 | 375 | 375.5 | 376 | 376.5 | 377 | 377.5 | 378 | 378.5 | 379 | 379.5 | 380 | 380.5 | 381 | 381.5 | 382 | 382.5 | 383 | 383.5 | 384 | 384.5 | 385 | 385.5 | 386 | 386.5 | 387 | 387.5 | 388 | 388.5 | 389 | 389.5 | 390 | 390.5 | 391 | 391.5 | 392 | 392.5 | 393 | 393.5 | 394 | 394.5 | 395 | 395.5 | 396 | 396.5 | 397 | 397.5 | 398 | 398.5 | 399 | 399.5 | 400 | 400.5 | 401 | 401.5 | 402 | 402.5 | 403 | 403.5 | 404 | 404.5 | 405 | 405.5 | 406 | 406.5 | 407 | 407.5 | 408 | 408.5 | 409 | 409.5 | 410 | 410.5 | 411 | 411.5 | 412 | 412.5 | 413 | 413.5 | 414 | 414.5 | 415 | 415.5 | 416 | 416.5 | 417 | 417.5 | 418 | 418.5 | 419 | 419.5 | 420 | 420.5 | 421 | 421.5 | 422 | 422.5 | 423 | 423.5 | 424 | 424.5 | 425 | 425.5 | 426 | 426.5 | 427 | 427.5 | 428 | 428.5 | 429 | 429.5 | 430 | 430.5 | 431 | 431.5 | 432 | 432.5 | 433 | 433.5 | 434 | 434.5 | 435 | 435.5 | 436 | 436.5 | 437 | 437.5 | 438 | 438.5 | 439 | 439.5 | 440 | 440.5 | 441 | 441.5 | 442 | 442.5 | 443 | 443.5 | 444 | 444.5 | 445 | 445.5 | 446 | 446.5 | 447 | 447.5 | 448 | 448.5 | 449 | 449.5 | 450 | 450.5 | 451 | 451.5 | 452 | 452.5 | 453 | 453.5 | 454 | 454.5 | 455 | 455.5 | 456 | 456.5 | 457 | 457.5 | 458 | 458.5 | 459 | 459.5 | 460 | 460.5 | 461 | 461.5 | 462 | 462.5 | 463 | 463.5 | 464 | 464.5 | 465 | 465.5 | 466 | 466.5 | 467 | 467.5 | 468 | 468.5 | 469 | 469.5 | 470 | 470.5 | 471 | 471.5 | 472 | 472.5 | 473 | 473.5 | 474 | 474.5 | 475 | 475.5 | 476 | 476.5 | 477 | 477.5 | 478 | 478.5 | 479 | 479.5 | 480 | 480.5 | 481 | 481.5 | 482 | 482.5 | 483 | 483.5 | 484 | 484.5 | 485 | 485.5 | 486 | 486.5 | 487 | 487.5 | 488 | 488.5 | 489 | 489.5 | 490 | 490.5 | 491 | 491.5 | 492 | 492.5 | 493 | 493.5 | 494 | 494.5 | 495 | 495.5 | 496 | 496.5 | 497 | 497.5 | 498 | 498.5 | 499 | 499.5 | 500 | 500.5 | 501 | 501.5 | 502 | 502.5 | 503 | 503.5 | 504 | 504.5 | 505 | 505.5 | 506 | 506.5 | 507 | 507.5 | 508 | 508.5 | 509 | 509.5 | 510 | 510.5 | 511 | 511.5 | 512 | 512.5 | 513 | 513.5 | 514 | 514.5 | 515 | 515.5 | 516 | 516.5 | 517 | 517.5 | 518 | 518.5 | 519 | 519.5 | 520 | 520.5 | 521 | 521.5 | 522 | 522.5 | 523 | 523.5 | 524 | 524.5 | 525 | 525.5 | 526 | 526.5 | 527 | 527.5 | 528 | 528.5 | 529 | 529.5 | 530 | 530.5 | 531 | 531.5 | 532 | 532.5 | 533 | 533.5 | 534 | 534.5 | 535 | 535.5 | 536 | 536.5 | 537 | 537.5 | 538 | 538.5 | 539 | 539.5 | 540 | 540.5 | 541 | 541.5 | 542 | 542.5 | 543 | 543.5 | 544 | 544.5 | 545 | 545.5 | 546 | 546.5 | 547 | 547.5 | 548 | 548.5 | 549 | 549.5 | 550 | 550.5 | 551 | 551.5 | 552 | 552.5 | 553 | 553.5 | 554 | 554.5 | 555 | 555.5 | 556 | 556.5 | 557 | 557.5 | 558 | 558.5 | 559 | 559.5 | 560 | 560.5 | 561 | 561.5 | 562 | 562.5 | 563 | 563.5 | 564 | 564.5 | 565 | 565.5 | 566 | 566.5 | 567 | 567.5 | 568 | 568.5 | 569 | 569.5 | 570 | 570.5 | 571 | 571.5 | 572 | 572.5 | 573 | 573.5 | 574 | 574.5 | 575 | 575.5 | 576 | 576.5 | 577 | 577.5 | 578 | 578.5 | 579 | 579.5 | 580 | 580.5 | 581 | 581.5 | 582 | 582.5 | 583 | 583.5 | 584 | 584.5 | 585 | 585.5 | 586 | 586.5 | 587 | 587.5 | 588 | 588.5 | 589 | 589.5 | 590 | 590.5 | 591 | 591.5 | 592 | 592.5 | 593 | 593.5 | 594 | 594.5 | 595 | 595.5 | 596 | 596.5 | 597 | 597.5 | 598 | 598.5 | 599 | 599.5 | 600 | 600.5 | 601 | 601.5 | 602 | 602.5 | 603 | 603.5 | 604 | 604.5 | 605 | 605.5 | 606 | 606.5 | 607 | 607.5 | 608 | 608.5 | 609 | 609.5 | 610 | 610.5 | 611 | 611.5 | 612 | 612.5 | 613 | 613.5 | 614 | 614.5 | 615 | 615.5 | 616 | 616.5 | 617 | 617.5 | 618 | 618.5 | 619 | 619.5 | 620 | 620.5 | 621 | 621.5 | 622 | 622.5 | 623 | 623.5 | 624 | 624.5 | 625 | 625.5 | 626 | 626.5 | 627 | 627.5 | 628 | 628.5 | 629 | 629.5 | 630 | 630.5 | 631 | 631.5 | 632 | 632.5 | 633 | 633.5 | 634 | 634.5 | 635 | 635.5 | 636 | 636.5 | 637 | 637.5 | 638 | 638.5 | 639 | 639.5 | 640 | 640.5 | 641 | 641.5 | 642 | 642.5 | 643 | 643.5 | 644 | 644.5 | 645 | 645.5 | 646 | 646.5 | 647 | 647.5 | 648 | 648.5 | 649 | 649.5 | 650 | 650.5 | 651 | 651.5 | 652 | 652.5 | 653 | 653.5 | 654 | 654.5 | 655 | 655.5 | 656 | 656.5 | 657 | 657.5 | 658 | 658.5 | 659 | 659.5 | 660 | 660.5 | 661 | 661.5 | 662 | 662.5 | 663 | 663.5 | 664 | 664.5 | 665 | 665.5 | 666 | 666.5 | 667 | 667.5 | 668 | 668.5 | 669 | 669.5 | 670 | 670.5 | 671 | 671.5 | 672 | 672.5 | 673 | 673.5 | 674 | 674.5 | 675 | 675.5 | 676 | 676.5 | 677 | 677.5 | 678 | 678.5 | 679 | 679.5 | 680 | 680.5 | 681 | 681.5 | 682 | 682.5 | 683 | 683.5 | 684 | 684.5 | 685 | 685.5 | 686 | 686.5 | 687 | 687.5 | 688 | 688.5 | 689 | 689.5 | 690 | 690.5 | 691 | 691.5 | 692 | 692.5 | 693 | 693.5 | 694 | 694.5 | 695 | 695.5 | 696 | 696.5 | 697 | 697.5 | 698 | 698.5 | 699 | 699.5 | 700 | 700.5 | 701 | 701.5 | 702 | 702.5 | 703 | 703.5 | 704 | 704.5 | 705 | 705.5 | 706 | 706.5 | 707 | 707.5 | 708 | 708.5 | 709 | 709.5 | 710 | 710.5 | 711 | 711.5 | 712 | 712.5 | 713 | 713.5 | 714 | 714.5 | 715 | 715.5 | 716 | 716.5 | 717 | 717.5 | 718 | 718.5 | 719 | 719.5 | 720 | 720.5 | 721 | 721.5 | 722 | 722.5 | 723 | 723.5 | 724 | 724.5 | 725 | 725.5 | 726 | 726.5 | 727 | 727.5 | 728 | 728.5 | 729 | 729.5 | 730 | 730.5 | 731 | 731.5 | 732 | 732.5 | 733 | 733.5 | 734 | 734.5 | 735 | 735.5 | 736 | 736.5 | 737 | 737.5 | 738 | 738.5 | 739 | 739.5 | 740 | 740.5 | 741 | 741.5 | 742 | 742.5 | 743 | 743.5 | 744 | 744.5 | 745 | 745.5 | 746 | 746.5 | 747 | 747.5 | 748 | 748.5 | 749 | 749.5 | 750 | 750.5 | 751 | 751.5 | 752 | 752.5 | 753 | 753.5 | 754 | 754.5 | 755 | 755.5 | 756 | 756.5 | 757 | 757.5 | 758 | 758.5 | 759 | 759.5 | 760 | 760.5 | 761 | 761.5 | 762 | 762.5 | 763 | 763.5 | 764 | 764.5 | 765 | 765.5 | 766 | 766.5 | 767 | 767.5 | 768 | 768.5 | 769 | 769.5 | 770 | 770.5 | 771 | 771.5 | 772 | 772.5 | 773 | 773.5 | 774 | 774.5 | 775 | 775.5 | 776 | 776.5 | 777 | 777.5 | 778 | 778.5 | 779 | 779.5 | 780 | 780.5 | 781 | 781.5 | 782 | 782.5 | 783 | 783.5 | 784 | 784.5 | 785 | 785.5 | 786 | 786.5 | 787 | 787.5 | 788 | 788.5 | 789 | 789.5 | 790 | 790.5 | 791 | 791.5 | 792 | 792.5 | 793 | 793.5 | 794 | 794.5 | 795 | 795.5 | 796 | 796.5 | 797 | 797.5 | 798 | 798.5 | 799 | 799.5 | 800 | 800.5 | 801 | 801.5 | 802 | 802.5 | 803 | 803.5 | 804 | 804.5 | 805 | 805.5 | 806 | 806.5 | 807 | 807.5 | 808 | 808.5 | 809 | 809.5 | 810 | 810.5 | 811 | 811.5 | 812 | 812.5 | 813 | 813.5 | 814 | 814.5 | 815 | 815.5 | 816 | 816.5 | 817 | 817.5 | 818 | 818.5 | 819 | 819.5 | 820 | 820.5 | 821 | 821.5 | 822 | 822.5 | 823 | 823.5 | 824 | 824.5 | 825 | 825.5 | 826 | 826.5 | 827 | 827.5 | 828 | 828.5 | 829 | 829.5 | 830 | 830.5 | 831 | 831.5 | 832 | 832.5 | 833 | 833.5 | 834 | 834.5 | 835 | 835.5 | 836 | 836.5 | 837 | 837.5 | 838 | 838.5 | 839 | 839.5 | 840 | 840.5 | 841 | 841.5 | 842 | 842.5 | 843 | 843.5 | 844 | 844.5 | 845 | 845.5 | 846 | 846.5 | 847 | 847.5 | 848 | 848.5 | 849 | 849.5 | 850 | 850.5 | 851 | 851.5 | 852 | 852.5 | 853 | 853.5 | 854 | 854.5 | 855 | 855.5 | 856 | 856.5 | 857 | 857.5 | 858 | 858.5 | 859 | 859.5 | 860 | 860.5 | 861 | 861.5 | 862 | 862.5 | 863 | 863.5 | 864 | 864.5 | 865 | 865.5 | 866 | 866.5 | 867 | 867.5 | 868 | 868.5 | 869 | 869.5 | 870 | 870.5 | 871 | 871.5 | 872 | 872.5 | 873 | 873.5 | 874 | 874.5 | 875 | 875.5 | 876 | 876.5 | 877 | 877.5 | 878 | 878.5 | 879 | 879.5 | 880 | 880.5 | 881 | 881.5 | 882 | 882.5 | 883 | 883.5 | 884 | 884.5 | 885 | 885.5 | 886 | 886.5 | 887 | 887.5 | 888 | 888.5 | 889 | 889.5 | 890 | 890.5 | 891 | 891.5 | 892 | 892.5 | 893 | 893.5 | 894 | 894.5 | 895 | 895.5 | 896 | 896.5 | 897 | 897.5 | 898 | 898.5 | 899 | 899.5 | 900 | 900.5 | 901 | 901.5 | 902 | 902.5 | 903 | 903.5 | 904 | 904.5 | 905 | 905.5 | 906 | 906.5 | 907 | 907.5 | 908 | 908.5 | 909 | 909.5 | 910 | 910.5 | 911 | 911.5 | 912 | 912.5 | 913 | 913.5 | 914 | 914.5 | 915 | 915.5 | 916 | 916.5 | 917 | 917.5 | 918 | 918.5 | 919 | 919.5 | 920 | 920.5 | 921 | 921.5 | 922 | 922.5 | 923 | 923.5 | 924 | 924.5 | 925 | 925.5 | 926 | 926.5 | 927 | 927.5 | 928 | 928.5 | 929 | 929.5 | 930 | 930.5 | 931 | 931.5 | 932 | 932.5 | 933 | 933.5 | 934 | 934.5 | 935 | 935.5 | 936 | 936.5 | 937 | 937.5 | 938 | 938.5 | 939 | 939.5 | 940 | 940.5 | 941 | 941.5 | 942 | 942.5 | 943 | 943.5 | 944 | 944.5 | 945 | 945.5 | 946 | 946.5 | 947 | 947.5 | 948 | 948.5 | 949 | 949.5 | 950 | 950.5 | 951 | 951.5 | 952 | 952.5 | 953 | 953.5 | 954 | 954.5 | 955 | 955.5 | 956 | 956.5 | 957 | 957.5 | 958 | 958.5 | 959 | 959.5 | 960 | 960.5 | 961 | 961.5 | 962 | 962.5 | 963 | 963.5 | 964 | 964.5 | 965 | 965.5 | 966 | 966.5 | 967 | 967.5 | 968 | 968.5 | 969 | 969.5 | 970 | 970.5 | 971 | 971.5 | 972 | 972.5 | 973 | 973.5 | 974 | 974.5 | 975 | 975.5 | 976 | 976.5 | 977 | 977.5 | 978 | 978.5 | 979 | 979.5 | 980 | 980.5 | 981 | 981.5 | 982 | 982.5 | 983 | 983.5 | 984 | 984.5 | 985 | 985.5 | 986 | 986.5 | 987 | 987.5 | 988 | 988.5 | 989 | 989.5 | 990 | 990.5 | 991 | 991.5 | 992 | 992.5 | 993 | 993.5 | 994 | 994.5 | 995 | 995.5 | 996 | 996.5 | 997 | 997.5 | 998 | 998.5 | 999 | 999.5 | 1000 | 1000.5 | 1001 | 1001.5 | 1002 | 1002.5 | 1003 | 1003.5 | 1004 | 1004.5 | 1005 | 1005.5 | 1006 | 1006.5 | 1007 | 1007.5 | 1008 | 1008.5 | 1009 | 1009.5 | 1010 | 1010.5 | 1011 | 1011.5 | 1012 | 1012.5 | 1013 | 1013.5 | 1014 | 1014.5 | 1015 | 1015.5 | 1016 | 1016.5 | 1017 | 1017.5 | 1018 | 1018.5 | 1019 | 1019.5 | 1020 | 1020.5 | 1021 | 1021.5 | 1022 | 1022.5 | 1023 | 1023.5 | 1024 | 1024.5 | 1025 | 1025.5 | 1026 | 1026.5 | 1027 | 1027.5 | 1028 | 1028.5 | 1029 | 1029.5 | 1030 | 1030.5 | 1031 | 1031.5 | 1032 | 1032.5 | 1033 | 1033.5 | 1034 | 1034.5 | 1035 | 1035.5 | 1036 | 1036.5 | 1037 | 1037.5 | 1038 | 1038.5 | 1039 | 1039.5 | 1040 | 1040.5 | 1041 | 1041.5 | 1042 | 1042.5 | 1043 | 1043.5 | 1044 | 1044.5 | 1045 | 1045.5 | 1046 | 1046.5 | 1047 | 1047.5 | 1048 | 1048.5 | 1049 | 1049.5 | 1050 | 1050.5 | 1051 | 1051.5 | 1052 | 1052.5 | 1053 | 1053.5 | 1054 | 1054.5 | 1055 | 1055.5 | 1056 | 1056.5 | 1057 | 1057.5 | 1058 | 1058.5 | 1059 | 1059.5 | 1060 | 1060.5 | 1061 | 1061.5 | 1062 | 1062.5 | 1063 | 1063.5 | 1064 | 1064.5 | 1065 | 1065.5 | 1066 | 1066.5 | 1067 | 1067.5 | 1068 | 1068.5 | 1069 | 1069.5 | 1070 | 1070.5 | 1071 | 1071.5 | 1072 | 1072.5 | 1073 | 1073.5 | 1074 | 1074.5 | 1075 | 1075.5 | 1076 | 1076.5 | 1077 | 1077.5 | 1078 | 1078.5 | 1079 | 1079.5 | 1080 | 1080.5 | 1081 | 1081.5 | 1082 | 1082.5 | 1083 | 1083.5 | 1084 | 1084.5 | 1085 | 1085.5 | 1086 | 1086.5 | 1087 | 1087.5 | 1088 | 1088.5 | 1089 | 1089.5 | 1090 | 1090.5 | 1091 | 1091.5 | 1092 | 1092.5 | 1093 | 1093.5 | 1094 | 1094.5 | 1095 | 1095.5 | 1096 | 1096.5 | 1097 | 1097.5 | 1098 | 1098.5 | 1099 | 1099.5 | 1100 | 1100.5 | 1101 | 1101.5 | 1102 | 1102.5 | 1103 | 1103.5 | 1104 | 1104.5 | 1105 | 1105.5 | 1106 | 1106.5 | 1107 | 1107.5 | 1108 | 1108.5 | 1109 | 1109.5 | 1110 | 1110.5 | 1111 | 1111.5 | 1112 | 1112.5 | 1113 | 1113.5 | 1114 | 1114.5 | 1115 | 1115.5 | 1116 | 1116.5 | 1117 | 1117.5 | 1118 | 1118.5 | 1119 | 1119.5 | 1120 | 1120.5 | 1121 | 1121.5 | 1122 | 1122.5 | 1123 | 1123.5 | 1124 | 1124.5 | 1125 | 1125.5 | 1126 | 1126.5 | 1127 | 1127.5 | 1128 | 1128.5 | 1129 | 1129.5 | 1130 | 1130.5 | 1131 | 1131.5 | 1132 | 1132.5 | 1133 | 1133.5 | 1134 | 1134.5 | 1135 | 1135.5 | 1136 | 1136.5 | 1137 | 1137.5 | 1138 | 1138.5 | 1139 | 1139.5 | 1140 | 1140.5 | 1141 | 1141.5 | 1142 | 1142.5 | 1143 | 1143.5 | 1144 | 1144.5 | 1145 | 1145.5 | 1146 | 1146.5 | 1147 | 1147.5 | 1148 | 1148.5 | 1149 | 1149.5 | 1150 | 1150.5 | 1151 | 1151.5 | 1152 | 1152.5 | 1153 | 1153.5 | 1154 | 1154.5 | 1155 | 1155.5 | 1156 | 1156.5 | 1157 | 1157.5 | 1158 | 1158.5 | 1159 | 1159.5 | 1160 | 1160.5 | 1161 | 1161.5 | 1162 | 1162.5 | 1163 | 1163.5 | 1164 | 1164.5 | 1165 | 1165.5 | 1166 | 1166.5 | 1167 | 1167.5 | 1168 | 1168.5 | 1169 | 1169.5 | 1170 | 1170.5 | 1171 | 1171.5 | 1172 | 1172.5 | 1173 | 1173.5 | 1174 | 1174.5 | 1175 | 1175.5 | 1176 | 1176.5 | 1177 | 1177.5 | 1178 | 1178.5 | 1179 | 1179.5 | 1180 | 1180.5 | 1181 | 1181.5 | 1182 | 1182.5 | 1183 | 1183.5 | 1184 | 1184.5 | 1185 | 1185.5 | 1186 | 1186.5 | 1187 | 1187.5 | 1188 | 1188.5 | 1189 | 1189.5 | 1190 | 1190.5 | 1191 | 1191.5 | 1192 | 1192.5 | 1193 | 1193.5 | 1194 | 1194.5 | 1195 | 1195.5 | 1196 | 1196.5 | 1197 | 1197.5 | 1198 | 1198.5 | 1199 | 1199.5 | 1200 | 1200.5 | 1201 | 1201.5 | 1202 | 1202.5 | 1203 | 1203.5 | 1204 | 1204.5 | 1205 | 1205.5 | 1206 | 1206.5 | 1207 | 1207.5 | 1208 | 1208.5 | 1209 | 1209.5 | 1210 | 1210.5 | 1211 | 1211.5 | 1212 | 1212.5 | 1213 | 1213.5 | 1214 | 1214.5 | 1215 | 1215.5 | 1216 | 1216.5 | 1217 | 1217.5 | 1218 | 1218.5 | 1219 | 1219.5 | 1220 | 1220.5 | 1221 | 1221.5 | 1222 | 1222.5 | 1223 | 1223.5 | 1224 | 1224.5 | 1225 | 1225.5 | 1226 | 1226.5 | 1227 | 1227.5 | 1228 | 1228.5 | 1229 | 1229.5 | 1230 | 1230.5 | 1231 | 1231.5 | 1232 | 1232.5 | 1233 | 1233.5 | 1234 | 1234.5 | 1235 | 1235.5 | 1236 | 1236.5 | 1237 | 1237.5 | 1238 | 1238.5 | 1239 | 1239.5 | 1240 | 1240.5 | 1241 | 1241.5 | 1242 | 1242.5 | 1243 | 1243.5 | 1244 | 1244.5 | 1245 | 1245.5 | 1246 | 1246.5 | 1247 | 1247.5 | 1248 | 1248.5 | 1249 | 1249.5 | 1250 | 1250.5 | 1251 | 1251.5 | 1252 | 1252.5 | 1253 | 1253.5 | 1254 | 1254.5 | 1255 | 1255.5 | 1256 | 1256.5 | 1257 | 1257.5 | 1258 | 1258.5 | 1259 | 1259.5 | 1260 | 1260.5 | 1261 | 1261.5 | 1262 | 1262.5 | 1263 | 1263.5 | 1264 | 1264.5 | 1265 | 1265.5 | 1266 | 1266.5 | 1267 | 1267.5 | 1268 | 1268.5 | 1269 | 1269.5 | 1270 | 1270.5 | 1271 | 1271.5 | 1272 | 1272.5 | 1273 | 1273.5 | 1274 | 1274.5 | 1275 | 1275.5 | 1276 | 1276.5 | 1277 | 1277.5 | 1278 | 1278.5 | 1279 | 1279.5 | 1280 | 1280.5 | 1281 | 1281.5 | 1282 | 1282.5 | 1283 | 1283.5 | 1284 | 1284.5 | 1285 | 1285.5 | 1286 | 1286.5 | 1287 | 1287.5 | 1288 | 1288.5 | 1289 | 1289.5 | 1290 | 1290.5 | 1291 | 1291.5 | 1292 | 1292.5 | 1293 | 1293.5 | 1294 | 1294.5 | 1295 | 1295.5 | 1296 | 1296.5 | 1297 | 1297.5 | 1298 | 1298.5 | 1299 | 1299.5 | 1300 | 1300.5 | 1301 | 1301.5 | 1302 | 1302.5 | 1303 | 1303.5 | 1304 | 1304.5 | 1305 | 1305.5 | 1306 | 1306.5 | 1307 | 1307.5 | 1308 | 1308.5 | 1309 | 1309.5 | 1310 | 1310.5 | 1311 | 1311.5 | 1312 | 1312.5 | 1313 | 1313.5 | 1314 | 1314.5 | 1315 | 1315.5 | 1316 | 1316.5 | 1317 | 1317.5 | 1318 | 1318.5 | 1319 | 1319.5 | 1320 | 1320.5 | 1321 | 1321.5 | 1322 | 1322.5 | 1323 | 1323.5 | 1324 | 1324.5 | 1325 | 1325.5 | 1326 | 1326.5 | 1327 | 1327.5 | 1328 | 1328.5 | 1329 | 1329.5 | 1330 | 1330.5 | 1331 | 1331.5 | 1332 | 1332.5 | 1333 | 1333.5 | 1334 | 1334.5 | 1335 | 1335.5 | 1336 | 1336.5 | 1337 | 1337.5 | 1338 | 1338.5 | 1339 | 1339.5 | 1340 | 1340.5 | 1341 | 1341.5 | 1342 | 1342.5 | 1343 | 1343.5 | 1344 | 1344.5 | 1345 | 1345.5 | 1346 | 1346.5 | 1347 | 1347.5 | 1348 | 1348.5 | 1349 | 1349.5 | 1350 | 1350.5 | 1351 | 1351.5 | 1352 | 1352.5 | 1353 | 1353.5 | 1354 | 1354.5 | 1355 | 1355.5 | 1356 | 1356.5 | 1357 | 1357.5 | 1358 | 1358.5 | 1359 | 1359.5 | 1360 | 1360.5 | 1361 | 1361.5 | 1362 | 1362.5 | 1363 | 1363.5 | 1364 | 1364.5 | 1365 | 1365.5 | 1366 | 1366.5 | 1367 | 1367.5 | 1368 | 1368.5 | 1369 | 1369.5 | 1370 | 1370.5 | 1371 | 1371.5 | 1372 | 1372.5 | 1373 | 1373.5 | 1374 | 1374.5 | 1375 | 1375.5 | 1376 | 1376.5 | 1377 | 1377.5 | 1378 | 1378.5 | 1379 | 1379.5 | 1380 | 1380.5 | 1381 | 1381.5 | 1382 | 1382.5 | 1383 | 1383.5 | 1384 | 1384.5 | 1385 | 1385.5 | 1386 | 1386.5 | 1387 | 1387.5 | 1388 | 1388.5 | 1389 | 1389.5 | 1390 | 1390.5 | 1391 | 1391.5 | 1392 | 1392.5 | 1393 | 1393.5 | 1394 | 1394.5 | 1395 | 1395.5 | 1396 | 1396.5 | 1397 | 1397.5 | 1398 | 1398.5 | 1399 | 1399.5 | 1400 | 1400.5 | 1401 | 1401.5 | 1402 | 1402.5 | 1403 | 1403.5 | 1404 | 1404.5 | 1405 | 1405.5 | 1406 | 1406.5 | 1407 | 1407.5 | 1408 | 1408.5 | 1409 | 1409.5 | 1410 | 1410.5 | 1411 | 1411.5 | 1412 | 1412.5 | 1413 | 1413.5 | 1414 | 1414.5 | 1415 | 1415.5 | 1416 | 1416.5 | 1417 | 1417.5 | 1418 | 1418.5 | 1419 | 1419.5 | 1420 | 1420.5 | 1421 | 1421.5 | 1422 | 1422.5 | 1423 | 1423.5 | 1424 | 1424.5 | 1425 | 1425.5 | 1426 | 1426.5 | 1427 | 1427.5 | 1428 | 1428.5 | 1429 | 1429.5 | 1430 | 1430.5 | 1431 | 1431.5 | 1432 | 1432.5 | 1433 | 1433.5 | 1434 | 1434.5 | 1435 | 1435.5 | 1436 | 1436.5 | 1437 | 1437.5 | 1438 | 1438.5 | 1439 | 1439.5 | 1440 | 1440.5 | 1441 | 1441.5 | 1442 | 1442.5 | 1443 | 1443.5 | 1444 | 1444.5 | 1445 | 1445.5 | 1446 | 1446.5 | 1447 | 1447.5 | 1448 | 1448.5 | 1449 | 1449.5 | 1450 | 1450.5 | 1451 | 1451.5 | 1452 | 1452.5 | 1453 | 1453.5 | 1454 | 1454.5 | 1455 | 1455.5 | 1456 | 1456.5 | 1457 | 1457.5 | 1458 | 1458.5 | 1459 | 1459.5 | 1460 | 1460.5 | 1461 | 1461.5 | 1462 | 1462.5 | 1463 | 1463.5 | 1464 | 1464.5 | 1465 | 1465.5 | 1466 | 1466.5 | 1467 | 1467.5 | 1468 | 1468.5 | 1469 | 1469.5 | 1470 | 1470.5 | 1471 | 1471.5 | 1472 | 1472.5 | 1473 | 1473.5 | 1474 | 1474.5 | 1475 | 1475.5 | 1476 | 1476.5 | 1477 | 1477.5 | 1478 | 1478.5 | 1479 | 1479.5 | 1480 | 1480.5 | 1481 | 1481.5 | 1482 | 1482.5 | 1483 | 1483.5 | 1484 | 1484.5 | 1485 | 1485.5 | 1486 | 1486.5 | 1487 | 1487.5 | 1488 | 1488.5 | 1489 | 1489.5 | 1490 | 1490.5 | 1491 | 1491.5 | 1492 | 1492.5 | 1493 | 1493.5 | 1494 | 1494.5 | 1495 | 1495.5 | 1496 | 1496.5 | 1497 | 1497.5 | 1498 | 1498.5 | 1499 | 1499.5 | 1500 | 1500.5 | 1501 | 1501.5 | 1502 | 1502.5 | 1503 | 1503.5 | 1504 | 1504.5 | 1505 | 1505.5 | 1506 | 1506.5 | 1507 | 1507.5 | 1508 | 1508.5 | 1509 | 1509.5 | 1510 | 1510.5 | 1511 | 1511.5 | 1512 | 1512.5 | 1513 | 1513.5 | 1514 | 1514.5 | 1515 | 1515.5 | 1516 | 1516.5 | 1517 | 1517.5 | 1518 | 1518.5 | 1519 | 1519.5 | 1520 | 1520.5 | 1521 | 1521.5 | 1522 | 1522.5 | 1523 | 1523.5 | 1524 | 1524.5 | 1525 | 1525.5 | 1526 | 1526.5 | 1527 | 1527.5 | 1528 | 1528.5 | 1529 | 1529.5 | 1530 | 1530.5 | 1531 | 1531.5 | 1532 | 1532.5 | 1533 | 1533.5 | 1534 | 1534.5 | 1535 | 1535.5 | 1536 | 1536.5 | 1537 | 1537.5 | 1538 | 1538.5 | 1539 | 1539.5 | 1540 | 1540.5 | 1541 | 1541.5 | 1542 | 1542.5 | 1543 | 1543.5 | 1544 | 1544.5 | 1545 | 1545.5 | 1546 | 1546.5 | 1547 | 1547.5 | 1548 | 1548.5 | 1549 | 1549.5 | 1550 | 1550.5 | 1551 | 1551.5 | 1552 | 1552.5 | 1553 | 1553.5 | 1554 | 1554.5 | 1555 | 1555.5 | 1556 | 1556.5 | 1557 | 1557.5 | 1558 | 1558.5 | 1559 | 1559.5 | 1560 | 1560.5 | 1561 | 1561.5 | 1562 | 1562.5 | 1563 | 1563.5 | 1564 | 1564.5 | 1565 | 1565.5 | 1566 | 1566.5 | 1567 | 1567.5 | 1568 | 1568.5 | 1569 | 1569.5 | 1570 | 1570.5 | 1571 | 1571.5 | 1572 | 1572.5 | 1573 | 1573.5 | 1574 | 1574.5 | 1575 | 1575.5 | 1576 | 1576.5 | 1577 | 1577.5 | 1578 | 1578.5 | 1579 | 1579.5 | 1580 | 1580.5 | 1581 | 1581.5 | 1582 | 1582.5 | 1583 | 1583.5 | 1584 | 1584.5 | 1585 | 1585.5 | 1586 | 1586.5 | 1587 | 1587.5 | 1588 | 1588.5 | 1589 | 1589.5 | 1590 | 1590.5 | 1591 | 1591.5 | 1592 | 1592.5 | 1593 | 1593.5 | 1594 | 1594.5 | 1595 | 1595.5 | 1596 | 1596.5 | 1597 | 1597.5 | 1598 | 1598.5 | 1599 | 1599.5 | 1600




Chart: 1st integral gap dependence
| 9 | 9.5 | 10 | 12 | 14 | 16 | 18 | 20 | 22 | 25 | 28 | 31 | 34 | 37 | 40 | 45 | 50 | 60 | 80 | 100 | 150 | 200 | 9 | 10 | 11 | 12 | 13 | 15 | 17 | 20 | 25 | 30 | 50 | 100 | 150 | 200
 | -0.02546 | -0.030981 | -0.020512 | -0.01406 | -0.0093059 | -0.021564 | -0.018039 | -0.0032273 | 0.01708 | 0.0079351 | 0.024845 | 0.045727 | 0.040433 | 0.05078
Chart: 1st integral gap dependence
| 9 | 9.5 | 10 | 12 | 14 | 16 | 18 | 20 | 22 | 25 | 28 | 31 | 34 | 37 | 40 | 45 | 50 | 60 | 80 | 100 | 150 | 200
 | 0.045278 | 0.065941 | 0.027084 | -0.023505 | -0.064882 | -0.080662 | -0.11194 | -0.09769 | -0.10432 | -0.081363 | -0.072049 | -0.055132 | -0.041261 | -0.026441 | -0.033327 | -0.013919 | -0.0036364 | 0.0061261 | 0.032489 | 0.039109 | 0.048432 | 0.052212
15µrad

## Slide 11: Pole tuning
Different pole movements create different Vertical or Horizontal kicks
Kick
Transverse coordinate

## Slide 12: Pole tuning results
Chart: Trajectory initial
Tmm2
Chart: Phase Error
Longitudinal coordinate
Deg.
Chart: Trajectory tuned
±5mm transversely
displaced trajectories
on-axis
Tmm2
(λU-filtered)
2.6 Tmm2 RMS

## Slide 13: Measurement Bench / Stretched Wire Probe
  Hall probe  vertical field 1D-probe Bell FWB710
  Pick-up coil  horizontal field ~5x10mm2, N~5000
  Stretched wire probe Single stretched wire (100µm CuBe)

## Slide 14: Hall probe calibration
DB~2µT
Field map of calibration magnet
Better alignment of Hall and NMR probe with motorized stages
Adjustment of pole face improves current dependent shift of field plateau

## Slide 15: Hall probe calibration correction by stretched wire
Wire
Hall
Hall flip
Gap,mm
I2, Tmm2
2nd Field integral gap dependence
before and after correction
Field integrals measured by stretched wire and hall probe, normal and 180°-flipped, were different: corresponding to ~50uT difference in the average field.
Hall probe calibration is represented by high order polynomial, built using NMR measurements in a calibration magnet; 10µT reproducibility (inclusive 180° flip)
By adjusting even order coefficients, it is possible to minimize difference between hall probe, and stretched wire.
Empirical correction for an error which is not yet understood
200 Tmm2

## Slide 16: Digital integrator’s noise level

## Slide 17: Budker Insitiute of Nuclear Physics: VsDC
 | VsDC2(CAN) | VsDC3(VME)
ADC | 2 ch, 300ksps, 24bit resolution | 
Input ranges | ±0.2V .. ±10V | ±0.2V; ±2V PGA
Integral SNR at 10 μs at 1 ms at 1 s | 5·10-5 10-6 5·10-7 | 10-5 5·10-7 10-7
Absolute error at 10 μs at 100 μs > 1 ms | 10-3 10-4 10-5 | 10-3 10-4 10-5
Non-linearity | 20 ppm max | 
Gain error | 5 ppm max | 
Offset error | ±1 V | 
Triggering accuracy | ±2 ns | 
From factor | 3U 4HP Eurocard | 6U 4HP Eurocard

## Slide 18: Low Noise Sigma-delta ADC Evaluation Boards
 | AD7195EBZ | ADS1232REF
Resolution, bits | 24 | 24
Output rate, Hz | 5 .. 4800 | 10 or 80
Input Range, PGA | ±20mV.. ±2.5V | ±40mV.. ±5V
Input Noise@10Hz | 11nV RMS | 17nV RMS
Offset error * | 0.5uV, 5nV/°C | 0.1uV, 10nV/°C
ADC Gain drift | 1ppm/°C | 2.5ppm/°C
Nonlinearity, max | 15ppm | 20ppm
Price | 59$ | 49$
Ideal Integrator 0.1s
Hz
*autocalibration

## Slide 19: Stretched wire continuous movement
Present setup:
Keithley 2182A DVM
1mm step motion
~400ms integration time
15 mTmm noise
1.5 sec per point
100 sec measurement (60pts)
Test setup:
AD7195EBZ
Continuous motion
10Hz rate, 10mm/s
3mTmm noise
0.1 sec per point
8 sec measurement

## Slide 20: Using single stretched wire to simulate various coils
Horizontal field 1st integral:
Two measurements required:
Noise √2 higher.
Derivative dI/dy measured:
Integration increases noise
No absolute value:
Additional measurements of absolute values required
Vertical field 2nd integral:

## Slide 21: Horizontal field 1st integral, linear background (2 points)

## Slide 22: Horizontal field 1st integral, parabolic background (3 points)

## Slide 23: Conclusion
FLASH II - IDs
  Improved mechanical design based on PETRA IDs
  Smaller end-kick gap dependence
  Transfer measurements consider the actual pole contour
  Installation of 12 IDs in late 2013
Magnetic Measurements
  Hall probe calibration issues observed
  Characterization of various integrators in order to improve wire & coil measurements