## Slide 1: Long Term Monitoring of the LCLS-I Undulators
Yurii Levashov for SLAC MMF group
IMMW 18, BNL, June 3-7 2013

## Slide 2: Outline
  LCLS -1 undulators
  Long-term magnetic measurements:
  Temperature control
  Hall probe calibration
  Reference magnet
  Reference undulator
  Results:
  Accumulated radiation dose
  Undulator re-measurement results

## Slide 3: LCLS-1 Undulator Hall
Undulator segments:
Fixed gap, hybrid
1.Undulator line tapered, 1%
Each segment tuned for
a specific position with
accuracy of 10-4 (1m gap).
2.Canted poles, 4.5mrad
Installed - 33
Spare - 6
Reference - 1
Total - 40
Tuning time ~ 1.5 years.
Location
S/N
Spare
undulators
LCLS – 1 operated since April 2009.
Number of photons > 1012 , Energy range 0.7keV ÷ 11.9 keV.

## Slide 4: Temperature measurements
Reference temperature probe,
(Hart Scientific probe 1504, calibrated in ice bath )
4-wire RTD
(Therm-X)
Aluminum block
Chart: Temperature sensor stability
| 39919 | 39965 | 41079 | 41134 | 41332 | 41410
Th.1 | 0.002 | -0.001 | 0.01 | 0.008 | 0.005 | 0.007
Th.2 | 0.004 | 0.007 | 0.009 | 0.006 | -0.004 | 0.008
Th.3 | 0.004 | 0 | 0.001 | 0.002 | 0.004 | 0.006
Th.4 | -0.003 | 0.006 | 0.005 | -0.002 | 0.005 | 0.007
Th.5 | -0.001 | 0.003 | 0.008 | -0.001 | 0.007 | 0.006
Temperature compensation built in the undulator design. Temperature K dependence measured – 0.00043/°C.
Temperature variations limited to ±2C due to possible undulator damage. In the tunnel – 0.5 C temperature
difference, end to end. Temperature corrections applied.
Tuned at constant room temperature 20±0.1°C.
5 RTD attached to the face, distributed along the undulator.
RTD calibration checked with reference probe.

## Slide 5: Hall probe measurements
Sentron XZM12, 5V/T, 2-D probe with HP3458A voltmeters.
Measurements at 8cm/sec, start and finish inside zero Gauss chambers, corrections for zero drift applied.
Triggering every 200μm, 2ms integration time.
The probe is re-calibrated periodically w.r.t. NMR in a calibration magnet.
Probe calibration checked in a reference magnet.

## Slide 6: Hall probe calibration
Chart: Field Uniformity
| -6 | -5 | -4 | -3 | -2 | -1 | 0 | 1 | 2 | 3 | 4 | 5 | 6
1.15kG | -0.67 | -0.24 | -0.12 | -0.07 | -0.05 | -0.02 | 0 | 0 | -0.01 | -0.02 | -0.06 | -0.14 | -0.45
2.26kG | -1.06 | -0.24 | -0.05 | -0.02 | -0.01 | 0 | 0 | -0.01 | -0.04 | -0.07 | -0.13 | -0.28 | -0.96
4.5kG | -1.8 | -0.25 | 0.1 | 0.1 | 0.07 | 0.03 | 0 | -0.04 | -0.09 | -0.14 | -0.23 | -0.54 | -1.56
6.7kG | -2.89 | -0.25 | 0.25 | 0.24 | 0.16 | 0.07 | 0 | -0.07 | -0.12 | -0.18 | -0.3 | -0.72 | -2.2
8.9kG | -0.23 | 0.37 | 0.34 | 0.22 | 0.11 | 0 | -0.1 | -0.17 | -0.26 | -0.45 | -1.05
11kG | -0.25 | 0.44 | 0.43 | 0.31 | 0.16 | 0 | -0.13 | -0.26 | -0.44 | -0.72 | -1.59
13kG | -1.92 | -0.51 | -0.07 | 0.09 | 0.11 | 0 | -0.19 | -0.51 | -1.02 | -1.72 | -3.18
14.8kG | -4.52 | -1.95 | -0.87 | -0.23 | 0.04 | 0 | -0.32 | -0.96 | -1.99 | -3.48 | -6.53
NMR Probe
Wide, 20 cm dia. poles
Metrolab PT2025 NMR probe (1 of 5)
Copper HP cooling block.
Probe calibrated at
constant temperature 20 °C

## Slide 7: Hall probe calibration, continue
Remove zero offset:
Fit polynomial of 10-th order:
Subtract linear part:
Calculate residuals:
Do calibration 2 times, compare,
take average.
Calibration accuracy – 0.2G.

## Slide 8: Hall probe gain drift
Chart: Probe gain change at -1.2T
| 39148 | 39174 | 39225 | 39238 | 39278 | 39514 | 39678 | 39728 | 39800 | 39951 | 40255 | 40336 | 40588 | 40606 | 40613 | 40738 | 40854 | 40991 | 41138 | 41288 | 41366
 | -17.2 | -15.73 | -17.52 | -15.2 | -15.51 | -15.35 | -13.74 | -14.62 | -14.62 | -16.28 | -20.85 | -19.7 | -17.79 | -17.88 | -20.74 | -19.08 | -20.47 | -20.12 | -18.2 | -20.97 | -18.96
Probe clamped in, never removed from G-10 holder.
Plastic screw
Probe
G-10 holder

## Slide 9: Reference magnet
Ceramic spacer
NMR probe
Temperature sensor
PMs
Iron yoke
Iron poles
Y
X
Z
Measured with every undulator:
1. Scan in Z, fit parabola, find maximum.
2. Scan in X, fit parabola, find maximum.
3. Scan in Y, fit parabola, find minimum.
4. Move into Zero-Gauss chamber,
measure and subtract probe zero offset.
5. Put NMR probe in, measure field.
6. Calculate difference Hall probe – NMR.
B = 0.36889T at 20°C

## Slide 10: Reference magnet, continue
Chart: Reference magnet measurements history
Probe re-calibrated | 0.68 | 0.68 | 0.68 | 0.68 | 0.68 | 0.68 | 0.68 | 0.68 | 0.68 | 0.68 | 0.68 | 0.68 | 0.68 | 0.68 | 0.68 | 0.68 | 0.68
±0.5 G
Measured for every undulator, before final dataset is taken.

## Slide 11: Fiducialization
Y.Levashov “LCLS Undulator Fiducialization.” Proc. IMMW14, Geneva, 2005
Hall probe positioning accuracy 2μm in Y and 6μm in X.
PM calibration ±5m.
CMM measurement ±5m.
Fiducialization accuracy:
5μm in Y,
10μm in X.
Tooling ball
Fiducialization magnets.
2 assemblies attached to each end of the undulator. Distances from center to tooling balls are known from calibration.
Hall Probe
MA
PM center
PM tooling ball
Undulator tooling ball

## Slide 12: Reference undulator
Chart: Date
| 39209 | 39240 | 39270 | 39332 | 39393 | 39423 | 39455 | 39486 | 39546 | 39576 | 39637 | 39668 | 39699 | 39729 | 39729 | 39760 | 39822 | 39853 | 39881 | 39912 | 39942 | 40003 | 40065 | 40095 | 40126 | 40156 | 40219 | 40278 | 40308 | 40339 | 40339 | 40431 | 40554 | 40613 | 40644 | 40674 | 40735 | 40827 | 40920 | 41133 | 41287 | 41346 | 41407
 | 3.4976 | 3.4974 | 3.4979 | 3.4981 | 3.4981 | 3.4981 | 3.4983 | 3.4984 | 3.4982 | 3.4985 | 3.4988 | 3.4984 | 3.4985 | 3.4986 | 3.4987 | 3.4988 | 3.4985 | 3.4985 | 3.4986 | 3.4986 | 3.4985 | 3.4986 | 3.4987 | 3.4984 | 3.4986 | 3.4985 | 3.4987 | 3.4988 | 3.4987 | 3.4991 | 3.4986 | 3.4986 | 3.4974 | 3.4969 | 3.4971 | 3.497 | 3.4973 | 3.4972 | 3.497 | 3.4969 | 3.4973 | 3.497
Production Run average | 3.4986 | 3.4986 | 3.4986 | 3.4986 | 3.4986 | 3.4986 | 3.4986 | 3.4986 | 3.4986 | 3.4986 | 3.4986 | 3.4986 | 3.4986 | 3.4986 | 3.4986 | 3.4986 | 3.4986 | 3.4986 | 3.4986 | 3.4986 | 3.4986 | 3.4986 | 3.4986 | 3.4986
New average | 3.4971 | 3.4971 | 3.4971 | 3.4971 | 3.4971 | 3.4971 | 3.4971 | 3.4971 | 3.4971 | 3.4971
Production run
Voltmeters re-calibrated
(∆K/K = 4.410-4)
A special undulator. Stays inside the Lab
Re-measured after 3 regular undulators and after Hall probe re-calibration.
Voltmeters are checked with Valhalla Scientific 2701C voltage standard.
* K is calculated at the same x position and corrected for temperature differences.

## Slide 13: Measurement of Radiation dose
Each segment has TLD detector on
upstream end, close to the beam.
Detector readings - every 4 months.
TLD
Max. radiation dose ~ 5rad/week.
Expected detectable radiation
damage in ~20 years.
Courtesy of H-D Nuhn (nuhn@slac.stanford.edu)

## Slide 14: Change in undulator strength
Courtesy of H-D Nuhn (nuhn@slac.stanford.edu)
50 times faster than predicted?!

## Slide 15: Results of magnetic measurements
Chart: Change in undulators’ strength (re-measurement - initially tuned)
| 40210 | 40210 | 40269 | 40269 | 40299 | 40299 | 40461 | 40483 | 40513 | 40603 | 40695 | 40725 | 40756 | 40858 | 41030 | 41122 | 41275 | 41276 | 41307 | 41307 | 41318 | 41336 | 41336 | 41365 | 41368 | 41405 | 41415
Undulator K | -3 | -6 | 5.6 | 7.6 | -2.3 | -2.3 | -6.6 | 2.3 | 2.4 | -5 | -1 | 1.8 | -1.6 | 14 | 5.1 | 1.8 | 1.1 | -12 | -12 | -12 | -0.6 | -7.8 | 4.5 | -7.8 | -15 | -12 | -15
Chart: ΔB/B·10-5
| 40210 | 40210 | 40269 | 40269 | 40299 | 40299 | 40461 | 40483 | 40513 | 40603 | 40695 | 40725 | 40756 | 40858 | 41030 | 41122 | 41275 | 41276 | 41307 | 41307 | 41318 | 41336 | 41336 | 41365 | 41368 | 41405 | 41415
Reference magnet | 1.626 | 2.439 | 18.6992 | 21.4092 | 22.4932 | 10.5691 | 3.794 | -4.336 | 5.4201 | 5.1491 | -4.065 | -1.897 | 1.355 | -0.271 | -6.7751 | -6.5041 | 11.6531 | -7.3171 | -15.7182 | -3.523 | -14.0921 | -1.355 | 4.336 | 2.168 | -13.0081 | -10.5691 | -13.5501
Total 25 undulators re-measured

## Slide 16: Summary
Long – term undulator monitoring program is in place for 4 years.
Had a problem with voltmeters. Have to apply correction to K measurements.
Accumulated radiation dose is small. Magnetic measurements show faster damage than predicted. Our measurement system is not as accurate as we thought. Error bars to be re-evaluated.
To early to report radiation damage.

## Slide 17: Thank you.